## Sheet1
Company/Manufacturer | BGM Instrument model | Assay target | Technology principle | Readout technology | Enzyme | Sample type | Sample site | Patient group | End user profile (training level) | Highest level of detection | Lowest level of detection | Repeatability CV | Calibration required by user | Sample volume | Time to result | Hematocrit range | Storage temp (device) | Operating humidity (device) | Storage conditions (strips/reagents) | Size (cm) | Weight (g) | Minimum operating temp | Maximum operating temp | Display | Results memory capacity (number of tests) | Type of connectivity | Battery type | Battery life (number of tests) | Strip ejector | Digital integration with applications or cloud based solutions | Regulatory approval/certification | Standards for product performance | Standards for the manufacturer | Organisation main website
Abbott | FreeStyle Freedom Lite | Blood glucose | Electrochemical | Coulometric | glucose oxidase (GOD) | venous blood, whole blood, capillary blood | upper arm, forearm, hand, fingertip, thigh, calf | adults, children | Self-administered | 27.8 mmol/L | 1.1 mmol/L | * | No | 0.3 μL | <1 min | 15-65% | -20 to 60°C | up to 90 % | <30°C | 8.4 x 5.1 x 1.6 | 45.4 | 4°C | 40°C | LCD | 400 | Company provided cable | 1x CR 2032 (3.0V coin cell battery) | 1000 | No | Yes | CE-IVDD, FDA 510k | * | ISO 13485, ISO 9001, MDSAP, Other | https://www.diabetescare.abbott/worldwide-locations.html
Abbott | Freestyle Lite | Blood glucose | Electrochemical | Coulometric | glucose oxidase (GOD) | capillary blood, whole blood, venous blood | fingertip, upper arm, palm | adults, children | Self-administered, Trained staff in health care facilities | 500 mg/dL | 20 mg/dL | * | No | 0.3 μL | <1 min | 15-65% | -20 to 60°C | up to 90 | <30°C | 7.4 x 4.0 x 1.7 | 31.2 | 4°C | 40°C | LCD | 400 | Company provided cable | 1x CR 2032 (3.0V coin cell battery) | 400 | No | Yes | FDA 510k | * | ISO 13485, ISO 9001, MDSAP, Other | https://www.diabetescare.abbott/worldwide-locations.html
Abbott | Freestyle Optium Neo H | Blood glucose, β-Ketones | * | Amperometric | glucose dehydrogenase - nicotinamide Adenine Dinucleotide (GDH-NAD) | capillary blood, whole blood, venous blood, arterial blood | fingertip, neonatal heel | adults, children, neonates | Self-administered,Trained staff in health care facilities | 27.8 mmol/L | 1.1 mmol/L | up to 4.3% | Yes | 0.6 μL | 5 sec | 15-65% | -20 to 60°C | up to 90 | <40°C | 8.68 x 5.97 x 0.87 | 37 | 10°C | 50°C | LCD | 1000 | USB | 2x CR 2032 (3.0V coin cell battery) | 3000 | No | Yes | CE-IVDD | ISO 15197:2013 | ISO 13485, ISO 9001, MDSAP, Other | https://www.diabetescare.abbott/worldwide-locations.html
Abbott | FreeStyle Precision Neo | Blood glucose | * | Amperometric | * | whole blood | fingertip, forearm, upper arm, palm | adults, children | Self-administered, Trained staff in health care facilities | 27.8 mmol/L | 1.1 mmol/L | * | No | 0.6 μL | 5 sec | 15-65% | -25 to 55°C | up to 90 | -20 to 60°C | 8.68 x 5.97 x 0.87 | 37 | 10°C | 50°C | touchscreen | 1000 | USB | 2x CR 2032 (3.0V coin cell battery) | 3000 | No | Yes | FDA 510k | ISO 15197:2013 | ISO 13485, ISO 9001, MDSAP, Other | https://www.diabetescare.abbott/worldwide-locations.html
ACON Laboratories, Inc. | On Call Express | Blood glucose | * | * | glucose oxidase (GOD) | capillary blood, whole blood | fingertip, palm, forearm | adults, children | Self-administered | 600 mg/dL | 20 mg/dL | * | No | 0.4 μL | 4 sec | 25-60% | * | up to 90 | 5 to 30°C | 8.5 x 4.9 x 1.64 | 50 | 5°C | 45°C | LCD | 300 | USB | 1x CR 2032 (3.0V coin cell battery) | * | Yes | Yes | FDA 510k | * | ISO 13485, FDA QSR | https://www.aconlabs.com/
ACON Laboratories, Inc. | On Call Express II | Blood glucose | * | * | glucose oxidase (GOD) | capillary blood, whole blood | fingertip, palm, forearm | adults, children | Self-administered | 600 mg/dL | 20 mg/dL | * | No | 0.4 μL | 4 sec | 25-60% | * | up to 90 | 5 to 30°C | 8.15 x 4.83 x 1.6 | 47.4 | 5°C | 45°C | LCD | 500 | USB | 2x CR 2032 (3.0V coin cell battery) | 1000 | Yes | Yes | FDA 510k | * | ISO 13485, FDA QSR | https://www.aconlabs.com/
ACON Laboratories, Inc. | On Call Pro | Blood glucose | * | * | glucose oxidase (GOD) | capillary blood, whole blood | forearm, fingertip, palm | adults, children | Trained staff in health care facilities | 600 mg/dL | 20 mg/dL | up to 2.7% | No | 0.4 μL | 5 sec | 30- 55% | * | up to 90 | 5 to 30°C | 8.5 x 4.9 x 1.64 | 50 | 5°C | 45°C | LCD | 300 | USB | 1x CR 2032 (3.0V coin cell battery) | 1000 | Yes | Yes | FDA 510k | * | ISO 13485, FDA QSR | https://www.aconlabs.com/
ACON Laboratories, Inc. | On Call Express Voice | Blood glucose | * | * | glucose oxidase (GOD) | capillary blood, whole blood | palm, forearm, fingertip | adults, children | Self-administered | 600 mg/dL | 20 mg/dL | * | No | 0.4 μL | 4 sec | 25-60% | * | up to 90 | 5 to 30°C | 10.0 x 6.08 x 2.39 | 90 | 5°C | 45°C | LCD | 500 | USB | 2x, AAA | 1000 | Yes | Yes | FDA 510k | * | ISO 13485, FDA QSR | https://www.aconlabs.com/
ACON Laboratories, Inc. | On Call Extra Voice | Blood glucose | Electrochemical | Biosensor | glucose oxidase (GOD) | capillary blood | * | No information | * | 600 mg/dL, 33.3 mmol/L | 10 mg/dL, 0.6 mmol/L | * | No | 0.4 μL | 4 sec | 25-60% | * | up to 90 | * | 10.0 x 6.1 x 2.4 | 47.4 | * | * | * | 500 | USB | 2x AAA batteries (1.5V ) | * | Yes | Yes | Unknown | * | ISO 13485, FDA QSR | https://www.aconlabs.com/
ACON Laboratories, Inc. | On Call Extra | Blood glucose | * | * | glucose oxidase (GOD) | capillary blood, whole blood | palm, fingertip, forearm | No information | Self-administered, Trained staff in health care facilities | * | * | * | No | 0.4 μL | 4 sec | 25-60% | * | up to 90 | 2 to 30°C | 8.1 x 4.8 x 1.6 | 90 | 5°C | 45°C | LCD | 500 | USB | 2x CR 2032 (3.0V coin cell battery) | 1000 | Yes | Yes | CE-IVDD | ISO 15197 2015. | ISO 13485, FDA QSR | https://www.aconlabs.com/
ACON Laboratories, Inc. | On Call Extra Mobile | Blood glucose | Electrochemical | Biosensor | glucose oxidase (GOD) | capillary blood, whole blood | fingertip, palm, forearm | No information | Self-administered, Trained staff in health care facilities | 33.3 mmol/L | 0.6 mmol/L | * | No | 0.4 μL | 4 sec | 25-60% | -20 to 50°C | up to 90 | 2 to 35°C | 7.5 x 4.0 x 1.3 | 36 | 5°C | 45°C | LCD | 500 | Bluetooth | 1x CR 2032 (3.0V coin cell battery) | 500 | No | Yes | Unknown | ISO 15197 2015. | ISO 13485, FDA QSR | https://www.aconlabs.com/
ACON Laboratories, Inc. | On Call GK Dual | Blood glucose | * | * | * | * | * | adults, children, neonates | Self-administered, Trained staff in health care facilities | 600 mg/dL | 20 mg/dL | * | No | 0.8 μL | 5 sec | * | * | up to 90 | 2 to 30°C | 9.03 x 5.8 x 1.94 | 64 | 5°C | 45°C | LCD | 450 | USB | 2x CR 2032 (3.0V coin cell battery) | 3000 | Yes | Yes | CE-IVDD | * | ISO 13485, FDA QSR | https://www.aconlabs.com/
ACON Laboratories, Inc. | On Call Sure | Blood glucose | * | * | glucose dehydrogenase - flavin adenin dinucleotid (GDH - FAD) | * | * | * | * | 600 mg/dL | 10 mg/dL | * | No | 0.6 μL | 5 sec | 10-70% | * | up to 90 | * | * | * | * | * | * | * | Bluetooth | * | * | * | Yes | CE-IVDD, FDA 510k | ISO 15197 2015. | ISO 13485, FDA QSR | https://www.aconlabs.com/
ACON Laboratories, Inc. | On Call Sure Sync | Blood glucose | * | * | glucose dehydrogenase - flavin adenin dinucleotid (GDH - FAD) | capillary blood, arterial blood, venous blood | neonatal heel, forearm, palm, fingertip | adults, children, neonates | Self-administered, Trained staff in health care facilities | 600 mg/dL | 10 mg/dL | * | No | 0.6 μL | 5 sec | 10-70% | -20 to 50°C | up to 90 | * | 9.0 x 6.0 x 1.6 | 72 | 5°C | 45°C | LCD | 1000 | Bluetooth | 2x CR 2032 (3.0V coin cell battery) | 3000 | Yes | Yes | CE-IVDD, FDA 510k | ISO 15197 2015. | ISO 13485, FDA QSR | https://www.aconlabs.com/
ACON Laboratories, Inc. | On Call Vivid | Blood glucose | * | * | * | capillary blood | fingertip, neonatal heel, palm, forearm | adults, neonates | Self-administered, Trained staff in health care facilities | 600 mg/dL | 10 mg/dL | * | No | 0.8 μL | 5 sec | 20-70% | * | up to 90 | 2 to 30°C | 8.96 x 5.8 x 2.17 | 60 | 5°C | 45°C | LCD | 500 | Company provided cable | 2x CR 2032 (3.0V coin cell battery) | 1000 | Yes | Yes | FDA 510k | * | ISO 13485, FDA QSR | https://www.aconlabs.com/
ACON Laboratories, Inc. | On Call Express Mobile | Blood glucose | * | * | glucose oxidase (GOD) | capillary blood, whole blood | fingertip, forearm, palm | * | * | 600 mg/dL | 20 mg/dL | * | No | 0.4 μL | 4 sec | 25-60% | * | up to 90 | * | 7.5 x 4.0 x 1.3 | * | 5°C | 45°C | * | 500 | Bluetooth | 1x CR 2032 (3.0V coin cell battery) | 500 | * | Yes | FDA 510k | * | ISO 13485, FDA QSR | https://www.aconlabs.com/
ARKRAY, Inc. | Shine Connex | Blood glucose | Electrochemical | * | * | capillary blood, whole blood | fingertip | adults, children | Self-administered | 600 mg/dL | 20 mg/dL | up to 4.3% | No | 0.5 μL | 5 sec | 15-65% | 0 to 50°C | up to 90 | Store test strips in a cool and dry place at a temperature between 34–86 °F (1–30 ºC) and 20–80 % relative humidity. Keep from direct sunlight. | 10.3 x 5.4 x 1.6 | 72.1 | 6°C | 44°C | LCD | 1000 | Bluetooth | 2x CR 2032 (3.0V coin cell battery) | 1000 | Yes | Yes | Unknown | * | ISO 13485, ISO 9001, Other | https://arkrayusa.com/
ARKRAY, Inc. | Shine Express | Blood glucose | Electrochemical | * | * | capillary blood, whole blood | fingertip | adults, children | Self-administered | 600 mg/dL | 20 mg/dL | up to 4.3% | No | 0.5 μL | 5 sec | 15-65% | 0 to 50°C | up to 90 | 1 to 30°c, 20-80% RH | 10.3 x 5.4 x 1.7 | 70.9 | 6°C | 44°C | LCD | 1000 | Company provided cable | 2x AAA batteries (1.5V ) | 3000 | Yes | Yes | Unknown | * | ISO 13485, ISO 9001, Other | https://arkrayusa.com/
ARKRAY, Inc. | Shine XL | Blood glucose | Electrochemical | * | * | capillary blood, whole blood | palm, forearm, thigh, calf, fingertip | adults, children | Self-administered | 600 mg/dL | 20 mg/dL | up to 4.1% | No | 0.5 μL | 5 sec | 20–60 % | 0 to 50°C | up to 90 | 1 to 30 ºC | 9.3 x 4.7 x 1.5 | 51.5 | 10°C | 40°C | LCD | 250 | Company provided cable | 2x CR 2032 (3.0V coin cell battery) | 5000 | No | Yes | Unknown | * | ISO 13485, ISO 9001, Other | https://arkrayusa.com/
ARKRAY, Inc. | Shine | Blood glucose | Electrochemical | * | * | capillary blood, whole blood | palm, forearm, thigh, calf, fingertip | adults, children | Self-administered | 600 mg/dL | 20 mg/dL | up to 4.3% | No | 0.5 μL | 5 sec | 20–60 % | 0 to 50°C | up to 90 | 1 to 30 ºC | 9.50 x 3.30 x 1.95 | 41.1 | 10°C | 40°C | LCD | 500 | Company provided cable | 2x CR 2032 (3.0V coin cell battery) | 1000 | No | Yes | Unknown | * | ISO 13485, ISO 9001, Other | https://arkrayusa.com/
ARKRAY, Inc. | Expression | Blood glucose | * | * | * | capillary blood, whole blood | fingertip, forearm, palm | adults, children | Self-administered | 600 mg/dL | 20 mg/dL | * | No | 0.8 μL | 6 sec | 30-55% | 4 to 50°C | up to 85 | 1 to 30 ºC | 9.3 x 5.8 x 2.05 | 79 | 10°C | 40°C | LCD | 300 | Company provided cable | 2x AAA batteries | 1000 | No | Yes | Unknown | * | ISO 13485, ISO 9001, Other | https://arkrayusa.com/
ARKRAY, Inc. | Vital | Blood glucose | Electrochemical | Biosensor | glucose oxidase (GOD) | capillary blood, whole blood | fingertip, palm | adults, children | Self-administered, Trained staff in health care facilities | 600 mg/dL | 20 mg/dL | * | No | 0.5 μL | 7 sec | 33–52% | 0 to 50°C | up to 80 | 4 to 30°C | 8.0 x 4.6 x 1.6 | 39 | 5°C | 45°C | LCD | 250 | Company provided cable | 1x CR 2032 (3.0V coin cell battery) | 2000+ | No | Yes | CE-IVDD, FDA 510k | * | ISO 13485, ISO 9001, Other | https://arkrayusa.com/
ARKRAY, Inc. | Assure platinum | Blood glucose | Electrochemical | Biosensor | glucose oxidase (GOD) | capillary blood, venous blood | fingertip | adults, children | Self-administered, Trained staff in health care facilities | 600 mg/dL | 20 mg/dL | up to 2.2% | No | 0.5 μL | 7 sec | 30–55% | 0 to 50°C | up to 80 | 4 to 30°C | 11.43 x 6.35 x 3.05 | 59 | 5°C | 45°C | LCD | 500 | Company provided cable | 2x AAA batteries (1.5V) | 5000+ | Yes | Yes | FDA 510k | ISO 15197 2003 | ISO 13485, ISO 9001, Other | https://arkrayusa.com/
ARKRAY, Inc. | Assure Prism multi | Blood glucose | Electrochemical | * | * | capillary blood, whole blood | palm, forearm, thigh, calf, fingertip | adults, children | Trained staff in health care facilities | 600 mg/dL | 20 mg/dL | up to 3.1% | No | 0.5 μL | 5 sec | 20-60% | 0 to 50°C | up to 90 | 4 to 30°C | 9.3 x 5.6 x 2.1 | 59.6 | 10°C | 45°C | LCD | 500 | Company provided cable | 2x CR 2032 (3.0V coin cell battery) | 1000 | Yes | Yes | FDA 510k | * | ISO 13485, ISO 9001, Other | https://arkrayusa.com/
ARKRAY, Inc. | S onyx GT - 7120 | Blood glucose | * | * | * | capillary blood, whole blood | * | * | * | 600 mg/dL | 10 mg/dL | * | * | 0.3 μL | 5 sec | 20-70% | * | up to 80 | * | 9.4 x 5.7 x 1.85 | 70 | 8°C | 40°C | * | 500 | USB, Bluetooth | * | * | * | Yes | CE-IVDD | ISO 15197 2015. | ISO 13485, ISO 9001, Other | https://arkrayusa.com/
ARKRAY, Inc. | S GT- 7110 | Blood glucose | * | * | * | capillary blood, whole blood | * | * | * | 600 mg/dL | 10 mg/dL | * | * | 0.3 μL | 5 sec | 20-70% | * | up to 80 | * | 8.4 x 5.0 x 1.76 | 47 | 8°C | 40°C | * | 500 | USB | * | * | Yes | Yes | CE-IVDD | ISO 15197 2015. | ISO 13485, ISO 9001, Other | https://arkrayusa.com/
ARKRAY, Inc. | W GT 7210 | Blood glucose | * | * | * | capillary blood, venous blood | * | * | * | 600 mg/dL | 10 mg/dL | * | * | 0.5 μL | 7 sec | 20-70% | * | up to 80 | * | 8.4 x 5.0 x 1.76 | 47 | 8°C | 40°C | * | 500 | USB | * | * | Yes | Yes | CE-IVDD | ISO 15197 2015. | ISO 13485, ISO 9001, Other | https://arkrayusa.com/
ARKRAY, Inc. | 01 - mini | Blood glucose | Electrochemical | Biosensor | glucose oxidase (GOD) | capillary blood, whole blood | fingertip, palm | adults, children | Self-administered, Trained staff in health care facilities | 600 mg/dL | 20 mg/dL | * | No | 0.3 μL | 7 sec | 30-54% | 0 to 50°C | * | 1 to 30°C | 9.2 x 3.0 x 1.4 | 31 | 10°C | 40°C | LCD | 50 | * | 1x CR 2032 (3.0V coin cell battery) or DL2032 | 3000 | No | * | FDA 510k | * | ISO 13485, ISO 9001, Other | https://arkrayusa.com/
ARKRAY, Inc. | 01 mini plus | Blood glucose | * | * | * | whole blood | * | * | * | 600 mg/dL | 10 mg/dL | * | * | 0.3 μL | 7 sec | * | * | up to 80 | * | 10.2 x 3.3 x 1.85 | 30 | 10°C | 40°C | LCD | 500 | * | 2x CR 2032 (3.0V coin cell battery) | * | Yes | Yes | Unknown | * | ISO 13485, ISO 9001, Other | https://arkrayusa.com/
ARKRAY, Inc. | Σ | Blood glucose | Electrochemical | Enzymatic assays | glucose oxidase (GOD) | whole blood | palm, fingertip | * | * | 600 mg/dL | 10 mg/dL | * | No | 0.5 μL | 7 sec | * | 0 to 50°C | up to 80 | 1 to 30°C | 8.3 x 4.7 x 1.5 | 39 | 9°C | 40°C | LCD | 250 | * | 1x CR 2032 (3.0V coin cell battery) | 2000 | Yes | Yes | CE-IVDD | * | ISO 13485, ISO 9001, Other | https://arkrayusa.com/
ARKRAY, Inc. | x mini plus gt 1960 | Blood glucose | * | * | * | capillary blood, whole blood | * | * | * | 600 mg/dL | 10 mg/dL | * | * | 0.3 μL | 5 sec | * | * | up to 80% | * | 10.2 x 3.3 x 1.85 | 46 | 10°C | 40°C | LCD | 500 | * | 2x CR 2032 (3.0V coin cell battery) | * | Yes | * | CE-IVDD | * | ISO 13485, ISO 9001, Other | https://arkrayusa.com/
ARKRAY, Inc. | Σ-mini | Blood glucose | Electrochemical | Enzymatic assays | glucose oxidase (GOD) | capillary blood, whole blood | * | * | * | 600 mg/dL | 10 mg/dL | * | No | 0.5 μL | 7 sec | * | * | up to 80 | * | 6.9 x 3.5 x 1.15 | 25 | 11°C | 40°C | LCD | 50 | * | 1x CR 2032 (3.0V coin cell battery) | * | No | * | Unknown | * | ISO 13485, ISO 9001, Other | https://arkrayusa.com/
ARKRAY, Inc. | G+ LINK | Blood glucose | * | * | * | * | * | * | * | * | * | * | No | * | * | * | * | * | * | * | * | * | * | * | * | Bluetooth | * | * | Yes | Yes | Unknown | * | ISO 13485, ISO 9001, Other | https://arkrayusa.com/
ARKRAY, Inc. | G+ | Blood glucose | * | * | * | * | * | * | * | * | * | * | No | 0.6 μL | 5.5 sec | * | 0 to 50°C | * | * | * | * | 10°C | 40°C | LCD | 450 | * | * | * | Yes | * | Unknown | ISO 15197:2013 | ISO 13485, ISO 9001, Other | https://arkrayusa.com/
ARKRAY, Inc. | X-METER | Blood glucose | Electrochemical | Biosensor | glucose dehydrogenase - flavin adenin dinucleotid (GDH - FAD) | capillary blood, neonatal blood | fingertip, forearm, upper arm | adults, children | Self-administered, Trained staff in health care facilities | 600 mg/dL | 20 mg/dL | * | No | 0.3 μL | 5 sec | 30-52% | 0 to 50°C | * | 1 to 30°C, < 80% RH | 5.08 x 9.91 x 1.27 | 30 | 10°C | 40°C | LCD | 360 | Company provided cable | 1x CR 2032 (3.0V coin cell battery) or DL2032 | ~2000 | No | Yes | CE-IVDD, FDA 510k | * | ISO 13485, ISO 9001, Other | https://arkrayusa.com/
Ascensia Diabetes Care Holdings AG | Contour | Blood glucose | * | * | * | capillary blood, whole blood | fingertip, forearm, palm, thigh, abdomen | * | * | 33.3 mmol/L | 0.6 mmol/L | up to 5.1% | No | 0.6 μL | 5 sec | * | * | up to 93 | * | 7.7 x 5.7 x 1.9 | 52.3 | 5°C | 45°C | * | 480 | Company provided cable | 2x CR 2032 (3.0V coin cell battery) or DL2032 | ~1000 | No | Yes | CE-IVDD | ISO 15197 2015. | * | https://www.ascensia.com/
Ascensia Diabetes Care Holdings AG | Contour care | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | up to 93 | * | * | * | * | * | * | * | * | * | * | * | * | CE-IVDD | ISO 15197:2013 | * | https://www.ascensia.com/
Ascensia Diabetes Care Holdings AG | Contour next | Blood glucose | * | * | * | venous blood, capillary blood, whole blood | * | adults, children, neonates | Self-administered, Trained staff in health care facilities | 33.3 mmol/L | 0.6 mmol/L | up to 2.1% | No | 0.6 μL | 5 sec | * | * | up to 93 | protect from humidity | 8.0 x 5.3 x 1.27 | 45 | 5°C | 45°C | LCD | 800 | USB | 2x CR 2032 (3.0V coin cell battery) or DL2032 | ~1000 | No | Yes | CE-IVDD | ISO 15197:2013, ISO 15197 2015. | * | https://www.ascensia.com/
Ascensia Diabetes Care Holdings AG | Contour next connected | Blood glucose | * | * | * | capillary blood, whole blood, venous blood | fingertip, palm | No information | Self-administered, Trained staff in health care facilities | 33.3 mmol/L | 0.6 mmol/L | up to 2.8% | No | 0.6 μL | 5 sec | 0-70% | -20 to 65°C | up to 93 | 0 to 30°C, 10–80% RH | * | 53 | 5°C | 45°C | LCD | 800 | USB | 2x CR 2032 (3.0V coin cell battery) or DL2032, 225mAh capacity | ~1000 | No | * | Unknown | ISO 15197:2013 | * | https://www.ascensia.com/
Ascensia Diabetes Care Holdings AG | Contour next ez | Blood glucose | Electrochemical | Amperometric | glucose dehydrogenase - flavin adenin dinucleotid (GDH - FAD) | capillary blood, whole blood, venous blood | fingertip | adults, children, neonates | Self-administered, Trained staff in health care facilities | 33.3 mmol/L | 0.6 mmol/L | up to 1.7% | No | 0.6 μL | 5 sec | * | * | up to 93 | * | 7.7 x 5.7 x 1.9 | 47.5 | 5°C | 45°C | LCD | 480 | Unknown | 2x CR 2032 (3.0V coin cell battery) or DL2032, 225mAh capacity | ~1000 | No | Yes | FDA 510k | ISO 15197:2013 | * | https://www.ascensia.com/
Ascensia Diabetes Care Holdings AG | Contour next gen | Blood glucose | Electrochemical | Amperometric | glucose dehydrogenase - flavin adenin dinucleotid (GDH - FAD) | capillary blood, whole blood | fingertip | adults, children | Self-administered | 600 mg/dL | 20 mg/dL | up to 2.8% | No | 0.6 μL | 5 sec | 20-50% | -20 to 60°C | up to 93 | 5 to 30°C, 10%–80% Relative Humidity (RH) | 7.85 x 5.6 x 1.8 | 53 | 5°C | 45°C | LCD | 800 | USB | 2x CR 2032 (3.0V coin cell battery) or DL2032, 225mAh capacity | ~1000 | No | Yes | FDA 510k, Health canada | ISO 15197:2013 | * | https://www.ascensia.com/
Ascensia Diabetes Care Holdings AG | Contour next link | Blood glucose | Electrochemical | Amperometric | glucose dehydrogenase - flavin adenin dinucleotid (GDH - FAD) | capillary blood, whole blood | fingertip, palm | adults, children | Self-administered | 33.3 mmol/L | 1.1 mmol/L | up to 2.4% | No | 0.6 μL | 5 sec | * | * | up to 93 | 0 to 30°C | 9.7 x 3.1 x 1.8 | 43 | 5°C | 45°C | LCD | 1000 | USB | Rechargeable battery | * | No | Yes | FDA 510k | ISO 15197:2013, ISO 15197 2015. | * | https://www.ascensia.com/
Ascensia Diabetes Care Holdings AG | Contour next link 2.4 | Blood glucose | Electrochemical | Amperometric | glucose dehydrogenase - flavin adenin dinucleotid (GDH - FAD) | capillary blood, whole blood | fingertip, palm | adults, children | Self-administered | 33.3 mmol/L | 1.1 mmol/L | up to 2.0% | No | 0.6 μL | 5 sec | * | * | up to 93 | 0 to 30°C | 9.7 x 3.1 x 1.8 | 43 | 5°C | 45°C | LCD | 1000 | * | Rechargeable battery | * | No | Yes | CE-IVDD | ISO 15197:2013, ISO 15197 2015. | * | https://www.ascensia.com/
Ascensia Diabetes Care Holdings AG | Contour next one | Blood glucose | Electrochemical | Amperometric | glucose dehydrogenase - flavin adenin dinucleotid (GDH - FAD) | capillary blood, whole blood | fingertip, palm | adults, children | Self-administered | 600 mg/dL | 20 mg/dL | up to 2.1% | No | 0.6 μL | 5 sec | * | * | up to 93 | 0 to 30°C, 10–80% RH | 9.7 x 2.8 x 1.49 | 36 | 5°C | 45°C | LCD | 800 | USB | 2x CR 2032 (3.0V coin cell battery) or DL2032, 225mAh capacity | 1000 | No | Yes | CE-IVDD, FDA 510k | ISO 15197:2013 | * | https://www.ascensia.com/
Ascensia Diabetes Care Holdings AG | Contour plus | Blood glucose | Electrochemical | Amperometric | glucose dehydrogenase - flavin adenin dinucleotid (GDH - FAD) | capillary blood, whole blood | fingertip, palm | adults, children | Self-administered,Trained staff in health care facilities | 33.3 mmol/L | 0.6 mmol/L | up to 2.8% | No | 0.6 μL | 5 sec | * | * | up to 93 | * | 7.7 x 5.7 x 1.9 | 47.5 | 5°C | 45°C | LCD | 480 | * | 2x CR 2032 (3.0V coin cell battery) or DL2032, 225mAh capacity | ~1000 | No | * | CE-IVDD | ISO 15197:2013 | * | https://www.ascensia.com/
Ascensia Diabetes Care Holdings AG | Contour plus blue | Blood glucose | Electrochemical | Amperometric | glucose dehydrogenase - flavin adenin dinucleotid (GDH - FAD) | venous blood, capillary blood, whole blood | * | adults, children | Self-administered, Trained staff in health care facilities | 33.3 mmol/L | 0.6 mmol/L | up to 1.9% | No | 0.6 μL | 5 sec | 0-70% | -20 to 65°C | up to 93 | 0 to 30°C, 10–80% RH | 7.85 x 5.6 x 1.8 | 53 | 5°C | 45°C | LCD | 800 | USB | 2x CR 2032 (3.0V coin cell battery) or DL2032, 225mAh capacity | ~1000 | No | Yes | CE-IVDD | ISO 15197:2013 | * | https://www.ascensia.com/
Ascensia Diabetes Care Holdings AG | Contour plus elite | Blood glucose | Electrochemical | Amperometric | glucose dehydrogenase - flavin adenin dinucleotid (GDH - FAD) | venous blood, capillary blood, whole blood | fingertip, palm | adults, children | Self-administered, Trained staff in health care facilities | 600 mg/dL | 10 mg/dL | up to 2.8% | No | 0.6 μL | 5 sec | 0% to 70% | -20 to 65°C | up to 93 | 0 to 30°C, 10–80% RH | 7.85 x 5.6 x 1.8 | 53 | 5°C | 45°C | LCD | 800 | USB | 2x CR 2032 (3.0V coin cell battery) or DL2032, 225mAh capacity | 1000 | No | Yes | CE-IVDD | ISO 15197:2013 | * | https://www.ascensia.com/
Ascensia Diabetes Care Holdings AG | Contour plus one | Blood glucose | Electrochemical | Amperometric | glucose dehydrogenase - flavin adenin dinucleotid (GDH - FAD) | venous blood, whole blood, capillary blood | fingertip, palm | adults, children | Self-administered, Trained staff in health care facilities | 33.3 mmol/L | 0.6 mmol/L | up to 2.3% | No | 0.6 μL | 5 sec | * | * | up to 93 | * | 9.7 x 2.8 x 1.49 | 36 | 5°C | 45°C | LCD | 800 | USB | 2x CR 2032 (3.0V coin cell battery) or DL2032, 225mAh capacity | 1000 | No | Yes | CE-IVDD | ISO 15197:2013 | * | https://www.ascensia.com/
Ascensia Diabetes Care Holdings AG | Contour ts | Blood glucose | Electrochemical | Amperometric | glucose dehydrogenase - flavin adenin dinucleotid (GDH - FAD) | capillary blood, whole blood, venous blood, arterial blood | fingertip | adults, children | Self-administered, Trained staff in health care facilities | 33.3 mmol/L | 0.6 mmol/L | up to 3.6% | No | 0.6 μL | 5 sec | * | * | up to 93 | * | 7.1 x 6.0 x 1.9 | 56.7 | 5°C | 45°C | LCD | 250 | USB | 1x CR 2032 (3.0V coin cell battery) or DL2032, 225mAh capacity | 1000 | No | Yes | CE-IVDD | ISO 15197:2013 | * | https://www.ascensia.com/
Ascensia Diabetes Care Holdings AG | Contour xt | Blood glucose | Electrochemical | Amperometric | glucose dehydrogenase - flavin adenin dinucleotid (GDH - FAD) | venous blood, whole blood, capillary blood | fingertip | adults, children | Self-administered, Trained staff in health care facilities | 33.3 mmol/L | 0.6 mmol/L | up to 1.9% | No | 0.6 μL | 5 sec | * | * | up to 93 | * | 7.7 x 5.7 x 1.9 | 47.5 | 5°C | 45°C | LCD | 480 | * | 2x CR 2032 (3.0V coin cell battery) or DL2032, 225mAh capacity | 1000 | No | Yes | CE-IVDD | ISO 15197:2013 | * | https://www.ascensia.com/
Bionime Corporation | Max Plus | Blood glucose | Electrochemical | Amperometric | glucose dehydrogenase (GDH) | capillary blood, whole blood | palm, forearm, fingertip, neonatal heel | adults, children, neonates | Self-administered, Trained staff in health care facilities | 33.3 mmol/L | 0.6 mmol/L | * | No | 0.75 μL | 5 sec | 10-70% | -10 to 60°C | up to 90 | * | 5.0 x 8.2 x 1.55 | 64 | 6°C | 44°C | LCD | 500 | * | 2x CR 2032 (3.0V coin cell battery) | 1000 | No | * | CE-IVDD, FDA 510k | ISO 15197:2013, ISO 15197 2015. | ISO 13485, Other | https://www.bionime.com/
Bionime Corporation | GM 700SB | Blood glucose | Electrochemical | Amperometric | glucose dehydrogenase (GDH) | capillary blood, whole blood, venous blood, arterial blood | fingertip, palm, forearm, neonatal heel | adults, children, neonates | Self-administered, Trained staff in health care facilities | 33.3 mmol/L | 0.6 mmol/L | * | No | 0.75 μL | 5 sec | 20-70% when bg < 200 mg/dl, 20-60% when bg >200 mg/dl | -10 to 60°C | up to 90 | 4 to 30°C, < 90% RH | 8.2 x 4.5 x 1.55 | 48 | 6°C | 44°C | LCD | 500 | Bluetooth | 1x CR 2032 (3.0V coin cell battery) | 1000 | No | Yes | CE-IVDD, FDA 510k | ISO 15197:2013, ISO 15197 2015. | ISO 13485, Other | https://www.bionime.com/
Bionime Corporation | Wiz | Blood glucose | Electrochemical | Amperometric | glucose oxidase (GOD) | capillary blood, whole blood | palm, upper arm, fingertip | adults, children | Self-administered, Trained staff in health care facilities | 33.3 mmol/L | 0.6 mmol/L | up to 2.8% | No | 0.75 μL | 5 sec | 20-60% | -10 to 60°C | up to 90 | 4 to 30°C, < 90% RH | 8.2 x 5.0 x 1.55 | 64 | 10°C | 40°C | LCD | 500 | * | 1x CR 2032 (3.0V coin cell battery) | 1000 | No | * | CE-IVDD, FDA 510k | ISO 15197:2013, ISO 15197 2015. | ISO 13485, Other | https://www.bionime.com/
Bionime Corporation | Wiz Plus | Blood glucose | Electrochemical | Amperometric | glucose oxidase (GOD) | capillary blood, whole blood | palm, upper arm, fingertip | adults, children | Self-administered, Trained staff in health care facilities | 33.3 mmol/L | 0.6 mmol/L | up to 2.8% | No | 0.75 μL | 5 sec | 20-60% | -10 to 60°C | up to 90 | 4 to 30°C, < 90% RH | 8.2 x 5.0 x 1.55 | 64 | 10°C | 40°C | LCD | 500 | Bluetooth | 2x CR 2032 (3.0V coin cell battery) | 1000 | No | Yes | CE-IVDD, FDA 510k | ISO 15197:2013, ISO 15197 2015. | ISO 13485, Other | https://www.bionime.com/
Bionime Corporation | Rightest GM 280 B | Blood glucose | Electrochemical | Amperometric | glucose oxidase (GOD) | capillary blood, whole blood | fingertip, palm, forearm | adults, children | Self-administered, Trained staff in health care facilities | 600 mg/dL | 10 mg/dL | <5% | No | 0.75 μL | 5 sec | 30-57% | -10 to 60°C | up to 90 | 4 to 30°C, < 90% RH | 8.2 x 4.5 x 1.55 | 48 | 10°C | 40°C | LCD | 500 | Bluetooth | 1x CR 2032 (3.0V coin cell battery) | 1000 | No | Yes | CE-IVDD, FDA 510k | ISO 15197:2013, ISO 15197 2015. | ISO 13485, Other | https://www.bionime.com/
Bionime Corporation | GE 333 | Blood glucose | Electrochemical | Amperometric | glucose oxidase (GOD) | capillary blood, whole blood | fingertip, palm, forearm | No information | * | 33.3 mmol/L | 0.6 mmol/L | * | No | 0.75 μL | 5 sec | 20-60% | -10 to 60°C | up to 90 | 4 to 30°C, 10%- 90% RH | 8.5 x 5.7 x 2.25 | 81 | 10°C | 40°C | LCD | 500 | Bluetooth | 2x AAA batteries (1.5V) | * | No | * | CE-IVDD | ISO 15197:2013, ISO 15197 2015. | ISO 13485, Other | https://www.bionime.com/
Bionime Corporation | Rightest GT 333 | Blood glucose | Electrochemical | Amperometric | glucose oxidase (GOD) | capillary blood, whole blood | fingertip, palm, forearm | adults, children | Self-administered, Trained staff in health care facilities | 600 mg/dL | 10 mg/dL | * | No | 0.75 μL | 5 sec | * | -10 to 60°C | up to 90 | 4 to 30°C, 10%- 90% RH | 8.5 x 5.7 x 2.25 | 86 | 10°C | 40°C | LCD | 500 | Bluetooth | 2x AAA batteries (1.5V) | 1000 | No | Yes | CE-IVDD, FDA 510k | * | * | https://www.bionime.com/
Bionime Corporation | GM 720 | Blood glucose | Electrochemical | Amperometric | glucose dehydrogenase - flavin adenin dinucleotid (GDH - FAD) | capillary blood, arterial blood, venous blood, whole blood | fingertip, neonatal heel | adults, children, neonates | Self-administered, Trained staff in health care facilities | 600 mg/dL | 10 mg/dL | < 5% | No | 0.7 μL | 5 sec | 20-70% for Glucose concentrations ≦ 200 mg/dL | -10 to 60°C | up to 90 | 4 to 30°C, 10%- 90% RH | 7.13 x 3.9 x 1.4 | 55 | 6°C | 44°C | LCD | 1000 | Unknown | 2x CR 2032 (3.0V coin cell battery) | 600 | No | * | CE-IVDD, FDA 510k | ISO 15197:2013, ISO 15197 2015. | ISO 13485, Other | https://www.bionime.com/
Bionime Corporation | Rightest GM 700 | Blood glucose | Electrochemical | Amperometric | glucose dehydrogenase (GDH) | capillary blood, arterial blood, venous blood, whole blood | fingertip, neonatal heel | adults, children, neonates | Self-administered, Trained staff in health care facilities | 600 mg/dL | 10 mg/dL | < 5% | No | 0.75 μL | 5 sec | 20-70% for Glucose concentrations ≦ 200 mg/dL | -10 to 60°C | up to 90 | 4 to 30°C, 10%- 90% RH | 9.6 x 4.6 x 1.75 | 62 | 6°C | 44°C | LCD | 1000 | Unknown | 1x CR 2032 (3.0V coin cell battery) | 1000 | No | * | CE-IVDD, FDA 510k | ISO 15197:2013, ISO 15197 2015. | ISO 13485, Other | https://www.bionime.com/
Bionime Corporation | GM 550 | Blood glucose | Electrochemical | Amperometric | glucose oxidase (GOD) | capillary blood, whole blood | fingertip, palm, forearm | adults, children | Self-administered | 600 mg/dL | 10 mg/dL | < 5% | No | 0.75 μL | 5 sec | 30-57% | -10 to 60°C | up to 90 | 4 to 30°C, 10%- 90% RH | 9.06 x 4.6 x 1.65 | 58 | 10°C | 40°C | LCD | 500 | Unknown | 2x CR 2032 (3.0V coin cell battery) | 1000 | No | * | CE-IVDD, FDA 510k | ISO 15197:2013, ISO 15197 2015. | ISO 13485, Other | https://www.bionime.com/
Bionime Corporation | Rightest GM 260 | Blood glucose | Electrochemical | Amperometric | glucose oxidase (GOD) | capillary blood, whole blood | fingertip, palm, forearm | adults, children | Self-administered | 600 mg/dL | 10 mg/dL | < 5% | No | 0.75 μL | 5 sec | 30-57% | -10 to 60°C | up to 90 | 4 to 30°C, 10%- 90% RH | 9.8 x 4.5 x 1.65 | 60 | 10°C | 40°C | LCD | 500 | Unknown | 1x CR 2032 (3.0V coin cell battery) | 1000 | No | * | CE-IVDD, FDA 510k, China NMPA | ISO 15197:2013, ISO 15197 2015. | ISO 13485, Other | https://www.bionime.com/
Bionime Corporation | Alpha | Blood glucose | Electrochemical | Amperometric | glucose oxidase (GOD) | capillary blood, whole blood | fingertip, palm, forearm | adults, children | Self-administered, Trained staff in health care facilities | 600 mg/dL | 10 mg/dL | < 5% | No | 0.75 μL | 5 sec | 30-57% | -10 to 60°C | up to 90 | 4 to 30°C, 10%- 90% RH | 9.06 x 4.60 x 1.65 | 58 | 10°C | 40°C | LCD | 500 | Unknown | 1x CR 2032 (3.0V coin cell battery) | 1000 | No | * | CE-IVDD | ISO 15197:2013, ISO 15197 2015. | ISO 13485, Other | https://www.bionime.com/
Bionime Corporation | Elsa | Blood glucose | Electrochemical | Amperometric | glucose oxidase (GOD) | capillary blood, whole blood | fingertip, palm, forearm | adults, children | Self-administered, Trained staff in health care facilities | 600 mg/dL | 10 mg/dL | < 5% | No | 0.75 μL | 5 sec | 30-57% | -10 to 60°C | up to 90 | 4 to 30°C, 10%- 90% RH | 9.5 x 4.38 x 1.3 | 48 | 10°C | 40°C | LCD | 500 | Unknown | 1x CR 2032 (3.0V coin cell battery) | 1000 | No | * | CE-IVDD | ISO 15197:2013, ISO 15197 2015. | ISO 13485, Other | https://www.bionime.com/
Bionime Corporation | GE Gluco Cpature - Cielo Series | Blood glucose | Electrochemical | Amperometric | glucose oxidase (GOD) | capillary blood, whole blood | fingertip, palm, forearm | * | * | 600 mg/dL | 10 mg/dL | * | No | 0.75 μL | * | 30-57% | -10 to 60°C | up to 90 | 4 to 30°C, 10%- 90% RH | 8.2 x 4.5 x 1.55 | 48 | 10°C | 40°C | LCD | 500 | Unknown | 1x CR 2032 (3.0V coin cell battery) | * | * | * | Unknown | * | ISO 13485, Other | https://www.bionime.com/
Bionime Corporation | GE Gluco Cpature Plus - Cielo Series | Blood glucose | Electrochemical | Amperometric | glucose oxidase (GOD) | capillary blood, whole blood | fingertip, palm, forearm | * | * | 600 mg/dL | 10 mg/dL | * | No | 0.75 μL | * | 30-57% | -10 to 60°C | up to 90 | 4 to 30°C, 10%- 90% RH | 8.2 x 4.5 x 1.55 | 48 | 10°C | 40°C | LCD | 500 | Bluetooth | 1x CR 2032 (3.0V coin cell battery) | * | * | Yes | Unknown | * | ISO 13485, Other | https://www.bionime.com/
Bionime Corporation | GE Gluco Capture - Oro Series | Blood glucose | Electrochemical | Amperometric | glucose dehydrogenase (GDH) | capillary blood, whole blood, venous blood, arterial blood | fingertip, palm, forearm, neonatal heel | * | * | 600 mg/dL | 10 mg/dL | * | No | 0.75 μL | * | 20-70% HCT range for Glucose concentrations ≦ 11.1 mmol/L (200 mg/dL) | -10 to 60°C | up to 90 | 4 to 30°C, 10%- 90% RH | 8.2 x 4.5 x 1.55 | 48 | 6°C | 44°C | LCD | 500 | Unknown | 1x CR 2032 (3.0V coin cell battery) | * | * | * | Unknown | * | ISO 13485, Other | https://www.bionime.com/
Bionime Corporation | GE Gluco Capture Plus - Oro Series | Blood glucose | Electrochemical | Amperometric | glucose dehydrogenase (GDH) | capillary blood, whole blood, venous blood, arterial blood | fingertip, palm, forearm, neonatal heel | * | * | 600 mg/dL | 10 mg/dL | * | No | 0.75 μL | * | 20-70% HCT range for Glucose concentrations ≦ 11.1 mmol/L (200 mg/dL) | -10 to 60°C | up to 90 | 4 to 30°C, 10%- 90% RH | 8.2 x 4.5 x 1.55 | 48 | 6°C | 44°C | LCD | 500 | Bluetooth | 1x CR 2032 (3.0V coin cell battery) | * | * | Yes | Unknown | * | * | https://www.bionime.com/
Bionime Corporation | GE 100 | Blood glucose | Electrochemical | Amperometric | glucose oxidase (GOD) | capillary blood, whole blood | fingertip, palm, forearm | * | * | 600 mg/dL | 10 mg/dL | * | No | 0.75 μL | * | 30-57% | -10 to 60°C | up to 90 | 4 to 30°C, 10%- 90% RH | 9.8 x 4.6 x 1.75 | 58 | 4°C | 30°C | LCD | 500 | Unknown | 1x CR 2032 (3.0V coin cell battery) | * | * | * | Unknown | * | ISO 13485, Other | https://www.bionime.com/
Bionime Corporation | GE 200 | Blood glucose | Electrochemical | Amperometric | glucose dehydrogenase (GDH) | capillary blood, whole blood, venous blood, arterial blood | fingertip, palm, forearm, neonatal heel | * | * | 600 mg/dL | 10 mg/dL | * | No | 0.75 μL | * | 20-70% for Glucose concentrations ≦ 11.1 mmol/L (200 mg/dL) | -10 to 60°C | up to 90 | 4 to 30°C, 10%- 90% RH | 9.8 x 4.6 x 1.75 | 70 | 4°C | 30°C | LCD | 1000 | Unknown | 2x CR 2032 (3.0V coin cell battery) | * | * | * | CE-IVDD, FDA 510k | * | ISO 13485, Other | https://www.bionime.com/
Bionime Corporation | Max Dual | Blood glucose, β-Ketones | Electrochemical | Amperometric | glucose dehydrogenase (GDH) | capillary blood, whole blood, venous blood, arterial blood | fingertip, palm, forearm, neonatal heel | * | * | 600 mg/dL | 10 mg/dL | * | No | 0.75 μL | * | 10-70% | -10 to 60°C | up to 90 | 4 to 30°C, 10%- 90% RH | 5.0 x 8.2 x 1.55 | 64 | 6°C | 44°C | LCD | 500 | Unknown | 1x CR 2032 (3.0V coin cell battery) | * | * | * | CE-IVDD | ISO 15197:2013, ISO 15197 2015. | ISO 13485, Other | https://www.bionime.com/
Bionime Corporation | Rightest GM 280 | Blood glucose | Electrochemical | Amperometric | glucose oxidase (GOD) | capillary blood, whole blood | Amperometric, fingertip, palm, forearm | adults, children | Self-administered, Trained staff in health care facilities | 600 mg/dL | 10 mg/dL | <5% | No | 0.75 μL | 5 sec | 30-57% | -10 to 60°C | up to 89 | 4 to 30°C, < 90% RH | 8.2 x 4.5 x 1.5 | 48 | 10°C | 40°C | LCD | 500 | Unknown | 1x CR 2032 (3.0V coin cell battery) | 1000 | No | * | CE-IVDD, FDA 510k | ISO 15197:2013, ISO 15197 2015. | ISO 13485, Other | https://www.bionime.com/
Bionime Corporation | GT 100 | Blood glucose | Electrochemical | Amperometric | glucose oxidase (GOD) | capillary blood, whole blood | fingertip, palm, forearm | adults, children | Self-administered, Trained staff in health care facilities | 600 mg/dL | 10 mg/dL | <5% | No | 0.75 μL | 5 sec | 30-57% | -10 to 60°C | up to 90 | 4 to 30°C, < 90% RH | 9.8 x 4.6 x 1.75 | 58 | 10°C | 40°C | LCD | 500 | Unknown | 1x CR 2032 (3.0V coin cell battery) | 1000 | No | * | CE-IVDD, FDA 510k | ISO 15197:2013 | ISO 13485, Other | https://www.bionime.com/
Dr. Morepen | GlucOne Blood Glucose Monitor Model BG 03 | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Unknown | * | ISO 9001, Other | https://www.morepen.com/
Enseval Medika Prima | Voice | Blood glucose | * | * | * | * | * | * | * | * | * | * | No | 0.5 μL | 5 sec | * | * | * | * | * | * | * | * | * | 5000 | * | * | * | * | * | Unknown | ISO 15197:2013 | ISO 13485, ISO 9001, Other | https://emp.co.id/about/
Enseval Medika Prima | Smart | Blood glucose | * | * | * | * | * | * | * | * | * | * | * | * | 5 sec | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Unknown | ISO 15197:2013 | ISO 13485, ISO 9001, Other | https://emp.co.id/about/
Enseval Medika Prima | On | Blood glucose | * | * | * | * | * | * | * | * | * | * | * | * | 5 sec | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Unknown | ISO 15197:2013 | ISO 13485, ISO 9001, Other | https://emp.co.id/about/
ForaCare, Inc. | 6 Connect | Blood glucose, β-Ketones | * | * | * | capillary blood, whole blood | fingertip | adults, children | Self-administered | 600 mg/dL | 10 mg/dL | * | No | 0.5 μL | * | 20-70% | -20 to 60°C | up to 85 | 2 to 30°C, 10-90% RH (non-condensing) | 8.98 x 5.49 x 1.8 | 46.1 | 8°C | 45°C | LCD | 1000 | Bluetooth | 1x AAA batteries (1.5V) | * | Yes | Yes | Unknown | ISO 15197:2013 | ISO 13485, Other | https://foracare.com/
ForaCare, Inc. | 6 plus | Blood glucose, β-Ketones , Hematocrit, Hemoglobin, total cholesterol, uric acid | * | * | * | capillary blood, whole blood, venous blood, arterial blood | neonatal heel | * | Self-administered, Trained staff in health care facilities | 600 mg/dL | 10 mg/dL | * | No | 0.5 μL | * | * | -20 to 60°C | up to 85 | 2 to 32°C, < 85% RH | 8.98 x 5.49 x 1.8 | 46.1 | 8°C | 45°C | LCD | 1000 | USB | 1x AAA batteries (1.5V) | * | Yes | Yes | Unknown | * | * | https://foracare.com/
ForaCare, Inc. | GTel | Blood glucose, β-Ketones | * | * | * | capillary blood, whole blood | fingertip | adults, children | Self-administered | 600 mg/dL | 10 mg/dL | * | No | * | * | * | -20 to 60°C | up to 90 | 2 to 30°C, 10% to 90% RH (non-condensing) | 9.8 x 5.5 x 1.55 | 63.2 | 8°C | 45°C | LCD | 1000 | * | 1x 3.7V Li-Ion rechargeable battery | * | Yes | * | FDA 510k | * | ISO 13485, Other | https://foracare.com/
ForaCare, Inc. | Premium V10 | Blood glucose | Electrochemical | Amperometric | glucose oxidase (GOD) | capillary blood, whole blood | fingertip, palm, forearm, upper arm | adults, children | Self-administered | 600 mg/dL | 20 mg/dL | * | No | 0.5 μL | * | * | -20 to 60°C | up to 85 | 4 to 40°C, < 85% RH | 8.87 x 5.56 x 2.31 | 48.8 | 10°C | 40°C | LCD | 450 | Company provided cable | 2x AAA batteries (1.5V) | * | Yes | Yes | Unknown | ISO 15197:2013, ISO 15197 2003 | ISO 13485, Other | https://foracare.com/
ForaCare, Inc. | Premium V10 BLE | Blood glucose | Electrochemical | Amperometric | glucose oxidase (GOD) | capillary blood, whole blood | fingertip, forearm, palm, upper arm | adults, children | Self-administered | 600 mg/dL | 20 mg/dL | * | No | 0.5 μL | * | * | -20 to 60°C | up to 85 | 4 to 40°C, < 85% RH | 8.87 x 5.56 x 2.31 | 48.8 | 10°C | 40°C | LCD | 450 | Bluetooth | 2x AAA batteries (1.5V) | * | Yes | Yes | Unknown | * | ISO 13485, Other | https://foracare.com/
ForaCare, Inc. | Test N'GO Advance | Blood glucose, β-Ketones | * | * | * | capillary blood, whole blood | fingertip | adults, children | Self-administered | 600 mg/dL | 10 mg/dL | * | No | 0.5 μL | * | 20-70% | -20 to 60°C | up to 85 | 2 to 30°C, 10% to 90% RH (non-condensing) | 8.98 x 5.49 x 1.8 | 46.1 | 8°C | 45°C | LCD | 1000 | Bluetooth | 1x AAA batteries (1.5V) | * | Yes | Yes | Unknown | * | ISO 13485, Other | https://foracare.com/
ForaCare, Inc. | Test N'GO Advance Voice | Blood glucose, β-Ketones | * | * | * | capillary blood, whole blood | fingertip | adults, children | Self-administered | 600 mg/dL | 10 mg/dL | * | No | 0.5 μL | * | 20-70% | -20 to 60°C | up to 85 | 2 to 30°C, 10% to 90% RH (non-condensing) | 9.05 x 5.51 x 2.09 | 48.4 | 8°C | 45°C | LCD | 1000 | Bluetooth | 2x AAA batteries (1.5V) | * | Yes | Yes | Unknown | * | ISO 13485, Other | https://foracare.com/
ForaCare, Inc. | TN'G | Blood glucose | * | * | * | capillary blood, whole blood | fingertip | adults, children | Self-administered | 600 mg/dL | 20 mg/dL | * | No | 0.5 μL | * | * | -20 to 60°C | up to 85 | * | 9.3 x 2.6 x 1.55 | 31 | 10°C | 40°C | LCD | 450 | Bluetooth | Li-Polymer battery | * | Yes | Yes | Unknown | * | ISO 13485, Other | https://foracare.com/
ForaCare, Inc. | TN'G Voice | Blood glucose | Electrochemical | Amperometric | glucose dehydrogenase (GDH) | capillary blood, whole blood | fingertip, palm, forearm, upper arm | adults, children | Self-administered | 600 mg/dL | 20 mg/dL | * | No | 0.5 μL | * | * | -20 to 60°C | up to 85 | 2 to 32°C, < 85% RH | 10.44 x 5.8 x 2.3 | 71.2 | 10°C | 40°C | LCD | 450 | Bluetooth | 2x AAA batteries (1.5V) | * | Yes | Yes | Unknown | * | ISO 13485, Other | https://foracare.com/
ForaCare, Inc. | V30 | Blood glucose | Electrochemical | Amperometric | glucose oxidase (GOD) | whole blood | fingertip, palm, upper arm, forearm, thigh, calf | adults, children | Self-administered | 600 mg/dL | 20 mg/dL | * | No | 0.5 μL | * | * | -20 to 60°C | up to 85 | 4 to 40°C, < 85% RH | 9.6 x 5.0 x 2.3 | 52 | 10°C | 40°C | LCD | 450 | Company provided cable | 2x AAA batteries (1.5V) | * | Yes | Yes | Unknown | ISO 15197 2003, ISO 15197:2013 | ISO 13485, Other | https://foracare.com/
ForaCare, Inc. | Pro Voice V9 | Blood glucose | Electrochemical | Amperometric | glucose oxidase (GOD) | capillary blood, whole blood | fingertip, palm, forearm, upper arm | adults, children | Self-administered | 600 mg/dL | 20 mg/dL | * | No | 0.7 μL | * | * | -20 to 60°C | up to 85 | 4 to 40°C, 10-85% RH | 9.6 x 4.6 x 2.0 | 68 | 10°C | 40°C | LCD | 450 | Unknown | 2x AAA batteries (1.5V) | * | No | * | Unknown | * | ISO 13485, Other | https://foracare.com/
ForaCare, Inc. | Pro Voice V8 | Blood glucose | Electrochemical | Amperometric | glucose oxidase (GOD) | capillary blood, whole blood | fingertip, palm, forearm, upper arm | adults, children | Self-administered | 600 mg/dL | 20 mg/dL | * | No | 0.75 μL | * | * | -20 to 60°C | up to 85 | 4 to 40°C, 10-85% RH | 9.6 x 4.5 x 2.3 | 76.2 | 10°C | 40°C | LCD | 450 | Unknown | 2x AAA batteries (1.5V) | * | No | * | Unknown | * | ISO 13485, Other | https://foracare.com/
ForaCare, Inc. | V12 | Blood glucose | Electrochemical | Amperometric | glucose oxidase (GOD) | capillary blood, whole blood | fingertip, palm, upper arm, forearm, thigh, calf | adults, children | Self-administered | 600 mg/dL | 20 mg/dL | * | No | 0.7 μL | * | * | -20 to 60°C | up to 85 | 4 to 40°C, < 85% RH | 9.6 x 4.5 x 2.3 | 76.15 | 10°C | 40°C | LCD | 450 | Unknown | 2x AAA batteries (1.5V) | * | No | Yes | Unknown | * | ISO 13485, Other | https://foracare.com/
Philosys/Gmate | On | Blood glucose | Electrochemical | Amperometric | * | capillary blood, whole blood | fingertip, forearm, upper arm, hand, thigh, calf | adults, children | Self-administered | 600 mg/dL | 10 mg/dL | 1.6-4.1% | No | 0.5 μL | 5 sec | 20-60% | 2 to 32°C | up to 90 | 2 to 32°C | 4.15 x 3.36 x 9.0 | 5.9 | 10°C | 40°C | * | No limit | * | No battery | * | No | Yes | CE-IVDD | * | * | https://www.philosys.com/
Philosys/Gmate | Smart | Blood glucose | Electrochemical | Amperometric | * | capillary blood, whole blood | fingertip, forearm, upper arm, thigh, calf, palm | adults, children | Self-administered | 600 mg/dL | 20 mg/dL | * | No | 0.5 μL | 5 sec | 20-60% | 2 to 32°C | up to 90 | 2 to 32°C | 2.1 x 4.27 x 8.8 | 4.2 | 10°C | 40°C | * | No limit | * | No battery | * | No | Yes | CE-IVDD, FDA 510k | * | * | https://www.philosys.com/
Philosys/Gmate | origin | Blood glucose | Electrochemical | Amperometric | * | * | fingertip, hand, calf, forearm, upper arm, thigh | adults, children | Self-administered, Trained staff in health care facilities | 600 mg/dL | 20 mg/dL | up to 4.4% | No | 0.5 μL | 5 sec | 20-60% | 2 to 32°C | up to 90 | 2 to 32°C | * | * | 10°C | 32°C | * | 500 | Unknown | 1x CR 2032 (3.0V coin cell battery) | * | No | Yes | CE-IVDD, FDA 510k | * | * | https://www.philosys.com/
Philosys/Gmate | Voice | Blood glucose | Electrochemical | Amperometric | * | capillary blood, whole blood | fingertip, hand, upper arm, forearm, thigh, calf | adults, children | Self-administered | 600 mg/dL | 20 mg/dL | * | No | 0.5 μL | 5 sec | 20-60% | 2 to 32°C | up to 90 | 2 to 32°C | 6.2 x 11.48 x 1.3 | 80 | 10°C | 40°C | LCD | 5000 | Unknown | 2x AAA batteries (1.5V) | * | No | Yes | CE-IVDD, FDA 510k | * | * | https://www.philosys.com/
Philosys/Gmate | Step | Blood glucose | Electrochemical | Amperometric | * | capillary blood, whole blood | fingertip, hand, upper arm, forearm, thigh, calf | children, adults | Self-administered, Trained staff in health care facilities | 600 mg/dL | 10 mg/dL | up to 3.5 | No | 0.5 μL | 5 sec | 20-60% | 2 to 32°C | up to 90 | 2 to 32°C | 3.67 x 61.7 x 1.58 | 30 | 10°C | 40°C | LCD | 500 | Unknown | 1x CR 2032 (3.0V coin cell battery) | * | No | No | CE-IVDD, FDA 510k | * | * | https://www.philosys.com/
Philosys/Gmate | Wheel | Blood glucose | Electrochemical | Amperometric | * | capillary blood, whole blood | fingertip, hand, upper arm, forearm, thigh, calf | children, adults | Self-administered, Trained staff in health care facilities | 600 mg/dL | 10 mg/dL | up to 3.1% | No | 0.5 μL | 5 sec | 20-60% | 2 to 32°C | * | 2 to 32°C | 5.6 x 8.8 x 2.36 | * | 10°C | 40°C | LCD | 500 | Unknown | 1x CR 2032 (3.0V coin cell battery) | * | Yes | Yes | CE-IVDD | * | * | https://www.philosys.com/
i-SENS, Inc. | CareSens N | Blood glucose | Electrochemical | Amperometric | glucose oxidase (GOD) | Fresh Capillary Whold Blood | fingertip | children, adults | Self-administered, Trained staff in health care facilities | 600 mg/dL | 20 mg/dL | up to 3.6% | No | 0.5 μL | 5 sec | 15-65% | 0 to 50°C | up to 90 | 1 to 30°C | 9.3 x 4.7 x 1.5 | 51.5 | 5°C | 50°C | LCD | 1000 | Unknown | 2x CR 2032 (3.0V coin cell battery) | 3000 | No | Yes | CE-IVDD, FDA 510k | * | ISO 13485 | www.i-sens.com
i-SENS, Inc. | CareSens N NFC | Blood glucose | Electrochemical | Amperometric | glucose oxidase (GOD) | Fresh Capillary Whold Blood | fingertip | children, adults | Self-administered, Trained staff in health care facilities | 600 mg/dL | 20 mg/dL | up to 3.2% | No | 0.5 μL | 5 sec | 15-65% | 0 to 50°C | up to 90 | 1 to 30°C | 9.3 x 4.7 x 1.5 | 51.5 | 5°C | 50°C | LCD | 500 | NFC | 1x CR 2032 (3.0V coin cell battery) | 2000 | No | Yes | CE-IVDD | ISO 15197:2013 | ISO 13485 | www.i-sens.com
i-SENS, Inc. | CareSens N Voice | Blood glucose | Electrochemical | Amperometric | glucose oxidase (GOD) | Fresh Capillary Whold Blood | fingertip | adults, children | Self-administered, Trained staff in health care facilities | 600 mg/dL | 20 mg/dL | up to 3.5% | No | 0.5 μL | 5 sec | 15-65% | 0 to 50°C | up to 90 | 1 to 30°C | 10.0 x 5.6 x 1.9 | 70 | 5°C | 50°C | LCD | 500 | Unknown | 2x AAA batteries (1.5V) | 1000 | Yes | Yes | CE-IVDD, FDA 510k | ISO 15197:2013 | ISO 13485 | www.i-sens.com
i-SENS, Inc. | CareSens N POP | Blood glucose | Electrochemical | Amperometric | glucose oxidase (GOD) | Fresh Capillary Whold Blood | fingertip | adults, children | Self-administered, Trained staff in health care facilities | 600 mg/dL | 20 mg/dL | up to 3.8% | No | 0.5 μL | 5 sec | 15-65% | 0 to 50°C | up to 90 | 1 to 30°C | 9.5 x 3.3 x 1.9 | 41.4 | 5°C | 50°C | LCD | 1000 | Unknown | 2x CR 2032 (3.0V coin cell battery) | 2000 | No | Yes | CE-IVDD, FDA 510k | * | ISO 13485 | www.i-sens.com
i-SENS, Inc. | CareSens N Premier | Blood glucose | Electrochemical | Amperometric | glucose oxidase (GOD) | Fresh Capillary Whold Blood | fingertip | adults, children | Self-administered, Trained staff in health care facilities | 600 mg/dL | 20 mg/dL | up to 4.4% | No | 0.5 μL | 5 sec | 15-65% | 0 to 50°C | up to 90 | 1 to 30°C | 10.3 x 5.4 x 1.7 | 71.7 | 5°C | 50°C | LCD | 1000 | Bluetooth | 2x CR 2032 (3.0V coin cell battery) | 1000 | Yes | Yes | CE-IVDD, FDA 510k | * | ISO 13485 | www.i-sens.com
i-SENS, Inc. | CareSens N Eco | Blood glucose | Electrochemical | Amperometric | glucose oxidase (GOD) | Fresh Capillary Whold Blood | fingertip | adults, children | Self-administered, Trained staff in health care facilities | 600 mg/dL | 20 mg/dL | up to 3% | No | 0.5 μL | 5 sec | 15-65% | 0 to 50°C | up to 90 | 1 to 30°C | 9.5 x 4.9 x 1.8 | 49 | 5°C | 45°C | LCD | 1000 | Unknown | 1x CR 2032 (3.0V coin cell battery) | 1000 | Yes | Yes | CE-IVDD | * | ISO 13485 | www.i-sens.com
i-SENS, Inc. | CareSens S Fit | Blood glucose | Electrochemical | Amperometric | glucose oxidase (GOD) | Fresh Capillary Whold Blood | fingertip | adults, children | Self-administered, Trained staff in health care facilities | 600 mg/dL | 20 mg/dL | up to 3% | No | 0.5 μL | 6 sec | 15-65% | 0 to 50°C | up to 90 | 1 to 30°C | 9.5 x 4.9 x 1.75 | 50.5 | 5°C | 45°C | LCD | 1000 | Unknown | 1x CR 2032 (3.0V coin cell battery) | 1000 | No | Yes | CE-IVDD, FDA 510k | * | ISO 13485 | www.i-sens.com
i-SENS, Inc. | CareSens H Beat | Blood glucose | Electrochemical | Amperometric | glucose dehydrogenase - flavin adenin dinucleotid (GDH - FAD) | capillary blood, whole blood, venous blood, arterial blood | fingertip | adults, children, neonates | Self-administered, Trained staff in health care facilities | 600 mg/dL | 10 mg/dL | up to 4.1% | No | 0.5 μL | 5 sec | 15-70% | 0 to 50°C | up to 90 | 1 to 30°C | 9.5 x 4.96 x 1.79 | 52.8 | 5°C | 45°C | LCD | 1000 | Unknown | 1x CR 2032 (3.0V coin cell battery) | 1000 | Yes | Yes | CE-IVDD | * | ISO 13485 | www.i-sens.com
i-SENS, Inc. | CareSens PRO | Blood glucose | Electrochemical | Amperometric | glucose dehydrogenase - flavin adenin dinucleotid (GDH - FAD) for BGM, HBDH for Ketone | capillary blood, venous blood, whole blood | fingertip | adults, children, neonates | Self-administered, Trained staff in health care facilities | 600 mg/dL | 20 mg/dL | up to 3.7% | No | 0.4 μL | 5 sec | 15-65% | 0 to 50°C | up to 90 | 1 to 30°C | 10.6 x 5.8 x 1.7 | 71.8 | 5°C | 45°C | LCD | 1000 | Unknown | 2x CR 2032 (3.0V coin cell battery) | 3000 | Yes | Yes | CE-IVDD | ISO 15197:2013 | ISO 13485 | www.i-sens.com
i-SENS, Inc. | CareSens Dual | Blood glucose, β-Ketones | Electrochemical | Amperometric | glucose dehydrogenase - flavin adenin dinucleotid (GDH - FAD) | capillary blood, venous blood, whole blood, arterial blood | fingertip | adults, children, neonates | Self-administered, Trained staff in health care facilities | 600 mg/dL | 10 mg/dL | up to 4.2% | No | 0.4 μL | 5 sec | 15-65% | 0 to 50°C | up to 90 | 1 to 30°C | 10.6 x 5.8 x 1.7 | 72.6 | 5°C | 45°C | LCD | 1000 | Unknown | 2x CR 2032 (3.0V coin cell battery) | 1000 | Yes | Yes | CE-IVDD | * | ISO 13485 | www.i-sens.com
i-SENS, Inc. | CareSens II Plus | Blood glucose | Electrochemical | Amperometric | glucose oxidase (GOD) | Fresh Capillary Whold Blood | fingertip | adults, children | Self-administered, Trained staff in health care facilities | 600 mg/dL | 20 mg/dL | up to 4.0% | Yes | 0.5 μL | 5 sec | 15-65% | 0 to 50°C | up to 90 | 1 to 30°C | 9.15 x 4.5 x 1.59 | 49.7 | 5°C | 45°C | LCD | 1000 | Unknown | 2x CR 2032 (3.0V coin cell battery) | 3000 | No | Yes | CE-IVDD | * | ISO 13485 | www.i-sens.com
i-SENS, Inc. | NoCoding 1 Plus | Blood glucose | Electrochemical | Amperometric | glucose oxidase (GOD) | Fresh Capillary Whold Blood | fingertip | adults, children | Self-administered, Trained staff in health care facilities | 600 mg/dL | 20 mg/dL | up to 4.1% | No | 0.5 μL | 5 sec | 15-65% | 0 to 50°C | up to 90 | 1 to 30°C | 10.84 x 5.58 x 1.5 | 65.5 | 5°C | 50°C | LCD | 1000 | Bluetooth | 2x CR 2032 (3.0V coin cell battery) | 1000 | No | Yes | CE-IVDD, FDA 510k | ISO 15197:2013 | ISO 13485 | www.i-sens.com
LifeScan IP Holdings | ultra plus flex | Blood glucose | Electrochemical | Amperometric | glucose dehydrogenase - flavin adenin dinucleotid (GDH - FAD) | capillary blood, whole blood | fingertip | adults, children | Self-administered | 600 mg/dL | 20 mg/dL | up to 2.13% | No | 0.4 μL | 5 sec | 20-60% | 5 to 30°C | up to 90 | 5 to 30°C, < 65% RH | 5.21 x 8.59 x 1.6 | 49.9 | 10°C | 40°C | LCD | 500 | USB, Bluetooth | 1x CR 2032 (3.0V coin cell battery) | * | No | Yes | FDA 510k | * | ISO 13486 | https://www.lifescan.com/
LifeScan IP Holdings | select plus | Blood glucose | Electrochemical | Biosensor | glucose oxidase (GOD) | capillary blood, whole blood | fingertip | adults, children | Self-administered, Trained staff in health care facilities | 33.3 mmol/L | 1.1 mmol/L | up to 3.59 | No | 1.0 μL | 5 sec | 30-50% | 5 to 30°C | up to 90 | 5 to 30°C | 4.3 x 10.1 x 1.56 | 200 | 10°C | 44°C | LCD | 500 | USB | 2x CR 2032 (3.0V coin cell battery) | * | No | Yes | CE-IVDD | ISO 15197:2013 | ISO 13487 | https://www.lifescan.com/
LifeScan IP Holdings | ultra 2 | Blood glucose | Electrochemical | Biosensor | glucose oxidase (GOD) | capillary blood, whole blood | fingertip | adults, children | Self-administered, Trained staff in health care facilities | 600 mg/dL | 20 mg/dL | * | Yes | * | 5 sec | 30-50% | * | up to 90 | * | 7.92 x 5.72 x 2.29 | 42.5 | 6°C | 43°C | LCD | 500 | USB | 1x CR 2032 (3.0V coin cell battery) | * | No | Yes | Unknown | * | ISO 13488 | https://www.lifescan.com/
LifeScan IP Holdings | verio flex | Blood glucose | Electrochemical | Amperometric | glucose dehydrogenase - flavin adenin dinucleotid (GDH - FAD) | capillary blood, whole blood | fingertip | adults, children | Self-administered | 600 mg/dL | 20 mg/dL | up to 2.26% | No | 0.4 μL | 5 sec | 20-60% | 5 to 30°C | up to 90 | 5 to 30°C, < 65% RH | 5.21 x 8.59 x 1.6 | 49.9 | 10°C | 40°C | LCD | 500 | USB, Bluetooth | 1x CR 2032 (3.0V coin cell battery) | * | No | Yes | CE-IVDD, FDA 510k | ISO 15197:2013 | ISO 13490 | https://www.lifescan.com/
LifeScan IP Holdings | ultra easy | Blood glucose | Electrochemical | Biosensor | glucose oxidase (GOD) | capillary blood, whole blood | fingertip, fingertip, palm, forearm | adults, children | Self-administered, Trained staff in health care facilities | 33.3 mmol/L | 1.1 mmol/L | * | No | * | 5 sec | 30-55% | <30°C | up to 90 | <30°C, Store each item in a cool, dry place | 10.8 x 3.20 x 1.70 | 40 | 6°C | 44°C | LCD | 500 | USB | 1x CR 2032 (3.0V coin cell battery) | * | No | Yes | CE-IVDD | * | ISO 13491 | https://www.lifescan.com/
LifeScan IP Holdings | verio | Blood glucose | Electrochemical | Amperometric | glucose dehydrogenase - flavin adenin dinucleotid (GDH - FAD) | capillary blood, whole blood | fingertip | adults, children | Self-administered | 600 mg/dL | 20 mg/dL | up to 2.44% | No | 0.4 μL | 5 sec | 20-60% | 5 to 30°C | up to 90 | 5 to 30°C | 2.51 x 5.18 x 8.00 | 85 | 10°C | 40°C | LCD | 500 | USB | 2x AAA batteries (1.5V) | * | No | Yes | FDA 510k | * | ISO 13492 | https://www.lifescan.com/
LifeScan IP Holdings | select plus flex | Blood glucose | Electrochemical | Biosensor | glucose oxidase (GOD) | capillary blood, whole blood | fingertip | adults, children | Self-administered, Trained staff in health care facilities | 600 mg/dL | 20 mg/dL | up to 3.58% | No | 1.0 μL | 5 sec | 30-55% | 5 to 30°C | up to 90 | 5 to 30°C | 5.2 x 8.6 x 1.6 | 50 | 10°C | 44°C | LCD | 500 | USB, Bluetooth | 1x CR 2032 (3.0V coin cell battery) | * | No | Yes | CE-IVDD | * | ISO 13493 | https://www.lifescan.com/
LifeScan IP Holdings | ultra plus reflect | Blood glucose | Electrochemical | Amperometric | glucose dehydrogenase - flavin adenin dinucleotid (GDH - FAD) | capillary blood, whole blood | fingertip | adults, children | Self-administered, Trained staff in health care facilities | 33.3 mmol/L | 1.1 mmol/L | up to 2.91% | No | 0.4 μL | 5 sec | 20-60% | 5 to 30°C | up to 90 | 5 to 30°C | 4.3 x 10.1 x 1.56 | < 53 | 6°C | 44°C | LCD | 750 | USB, Bluetooth | 2x CR 2032 (3.0V coin cell battery) | * | No | Yes | CE-IVDD | * | ISO 13494 | https://www.lifescan.com/
LifeScan IP Holdings | verio reflect | Blood glucose | Electrochemical | Amperometric | glucose dehydrogenase - flavin adenin dinucleotid (GDH - FAD) | capillary blood, whole blood | fingertip | adults, children | Self-administered | 600 mg/dL | 20 mg/dL | up to 3.07 % | No | 0.4 μL | 5 sec | 20-60% | 5 to 30°C | up to 90 | 5 to 30°C, < 90% RH | 4.29 x 10.1 x 1.55 | < 57 | 6°C | 43°C | LCD | 750 | USB, Bluetooth | 2x CR 2032 (3.0V coin cell battery) | * | No | Yes | FDA 510k | * | ISO 13496 | https://www.lifescan.com/
Empecs | Medisign MM1100 | Blood glucose | Electrochemical | Amperometric | glucose oxidase (GOD) | capillary blood, whole blood | fingertip, palm, forearm | adults, children | Self-administered | 600 mg/dL | 20 mg/dL | * | No | 0.5 μL | 5 sec | 30-55% | 4 to 30°C | up to 90 | 4 to 30°C, 10-90% RH | 1.8 x 5.1 x 8.8 | 48 | 10°C | 40°C | LCD | 300 | USB, Bluetooth | 2x CR 2032 (3.0V coin cell battery) | 1000 | Yes | Yes | FDA 510k | * | ISO 13485 | https://empecs.com/
Ypsomed AG | Unio Cara | Blood glucose | Electrochemical | Amperometric | glucose dehydrogenase - flavin adenin dinucleotid (GDH - FAD) | capillary blood, whole blood, venous blood, arterial blood | fingertip, palm, forearm | adults, children | Self-administered, Trained staff in health care facilities | 600 mg/dL | 10 mg/dL | * | No | 0.7 μL | 5 sec | 20–70 % for glucose ≤ 200 mg/dL, 20–60 % for glucose > 200 mg/dL | -10 to 60°C | up to 90 | 4 to 30 °C, < 90 % RH | 8.2 x 5.0 x 1.55 | 64 | 6°C | 44°C | LCD | 1000 | Bluetooth, USB | 2x CR 2032 (3.0V coin cell battery) | 1000 | No | Yes | CE-IVDD | ISO 15197:2013 | ISO 13485, MDSAP, Other | https://www.ypsomed.com/en/
Ypsomed AG | Aveo | Blood glucose | Electrochemical | Amperometric | glucose dehydrogenase - flavin adenin dinucleotid (GDH - FAD) | capillary blood, whole blood, venous blood | neonatal heel, fingertip | adults, children, neonates | Self-administered, Trained staff in health care facilities | 33.3 mmol/L | 0.6 mmol/L | up to 4.1% | No | 0.5 μL | 5 sec | 15-70% | 0 to 50°C | up to 90 | 1 to 30°C | 10.3 x 5.4 x 1.5 | 67 | 5°C | 45°C | LCD | 1000 | Bluetooth, USB | 2x CR 2032 (3.0V coin cell battery) | 1000 | Yes | Yes | CE-IVDD | ISO 15197:2013 | ISO 13485, MDSAP, Other | https://www.ypsomed.com/en/
Ypsomed AG | Unio Neva | Blood glucose | Electrochemical | Amperometric | glucose dehydrogenase - flavin adenin dinucleotid (GDH - FAD) | capillary blood, whole blood, venous blood, arterial blood | fingertip, palm, forearm | adults, children | Self-administered, Trained staff in health care facilities | 33.3 mmol/L | 0.6 mmol/L | * | No | 0.7 μL | 5 sec | 20–70 % for glucose ≤ 11.1 mmol/L, 20–60 % for glucose > 11.1 mmol/L | -10 to 60°C | up to 90 | 4 to 30 °C, < 90 % RH | 7.15 x 3.95 x 1.4 | 55 | 6°C | 44°C | LCD | 1000 | Bluetooth, USB | 2x CR 2032 (3.0V coin cell battery) | 600 | No | Yes | CE-IVDD | ISO 15197:2013 | ISO 13485, MDSAP, Other | https://www.ypsomed.com/en/
Ypsomed AG | UnioPura och mylife Pura X | Blood glucose | Electrochemical | Biosensor | glucose oxidase (GOD) | capillary blood, whole blood | fingertip, forearm, palm | adults, children | Self-administered, Trained staff in health care facilities | 33.3 mmol/L | 0.6 mmol/L | * | No | 0.75 μL | 5 sec | 30-57% | -10 to 60°C | up to 90 | 4 to 30 °C, < 90 % RH | 9.06 x 4.6 x 1.65 | 58 | 10°C | 40°C | LCD | 500 | Unknown | 2x CR 2032 (3.0V coin cell battery) | 1000 | No | Yes | Unknown | ISO 15197:2013 | ISO 13485, MDSAP, Other | https://www.ypsomed.com/en/
Ypsomed AG | Unio | Blood glucose | Electrochemical | Amperometric | glucose dehydrogenase - flavin adenin dinucleotid (GDH - FAD) | capillary blood, arterial blood, venous blood, whole blood | fingertip, palm, forearm | adults, children | * | 600 mg/dL | 10 mg/dL | up to 2.0% | No | 0.7 μL | 5 sec | 20–70 % for glucose ≤ 200 mg/dL (11.1 mmol/L), 20–60 % for glucose > 200 mg/dL (11.1 mmol/L) | * | * | * | * | * | 6°C | 44°C | LCD | 1000 | USB | 2x CR 2032 (3.0V coin cell battery) | 600 | No | Yes | Unknown | ISO 15197:2013 | ISO 13485, MDSAP, Other | https://www.ypsomed.com/en/
F. Hoffmann-La Roche Ltd (Roche) | Accu Chek Active Model (GB) | Blood glucose | * | Photometric | glucose dehydrogenase - flavin adenin dinucleotid (GDH - FAD) | * | fingertip and alternative site testing | adults, children | Self-administered, Trained staff in health care facilities | 600 mg/dL | 10 mg/dL | 1.8% to 2.2% | No | 1.0 μL - 2.0 μL | 5 sec | 20-70 % | -25 to 70°C (without batteries), -25 to 50°C (with batteries) | up to 85 | 2 to 30°C | 9.78 x 4.68 x 1.91 | 50 | 8°C | 45°C | LCD | 500 | USB | 1x CR 2032 (3.0V coin cell battery) | 1000 | No | Yes | CE-IVDD | ISO 15197:2013-05 EN ISO 15197:2015-06 | ISO 13485, Other | https://www.roche.com/
F. Hoffmann-La Roche Ltd (Roche) | Accu Chek Guide | Blood glucose | Electrochemical | Amperometric | glucose dehydrogenase - flavin adenin dinucleotid (GDH - FAD) | capillary blood, whole blood | fingertip, palm, upper arm | adults, children | Self-administered | 600 mg/dL | 20 mg/dL | up to 3.5% | No | 0.6 μL | 4 sec | 10-65% | -25 to 70°C | * | 4 to 30°C, 10-90% RH | 8.0 x 4.7 x 2.0 | 40 | * | * | LCD | 752 | USB, Bluetooth | 2x CR 2032 (3.0V coin cell battery) | * | Yes | Yes | CE-IVDD, FDA 510k | ISO 15197:2013 | ISO 13485, Other | https://www.roche.com/
F. Hoffmann-La Roche Ltd (Roche) | Accu Chek Guide Me | Blood glucose | * | * | * | capillary blood, whole blood | fingertip, palm, upper arm | adults, children | Self-administered | 600 mg/dL | 20 mg/dL | * | No | * | * | 10-65% | 4 to 45°C | * | * | 7.6 x 4.8 x 1.6 | 43 | * | * | LCD | 750 | USB, Bluetooth | 2x CR 2032 (3.0V coin cell battery) | * | No | Yes | FDA 510k | ISO 15197 2015. | ISO 13485, Other | https://www.roche.com/
F. Hoffmann-La Roche Ltd (Roche) | Accu Chek Instant | Blood glucose | Electrochemical | Amperometric | glucose dehydrogenase - flavin adenin dinucleotid (GDH - FAD) | capillary blood, whole blood, venous blood, arterial blood | fingertip, palm, upper arm, forearm, neonatal heel | adults, children, neonates | Self-administered, Trained staff in health care facilities | 33.3 mmol/L | 0.6 mmol/L | * | No | 0.6 μL | 4 sec | 10-65% | 4 to 45°C | * | 4 to 45°C | 7.71 x 4.86 x 1.53 | * | 4 | 45°C | LCD | 720 | USB, Bluetooth | 2x CR 2032 (3.0V coin cell battery) | * | Yes | Yes | CE-IVDD | ISO 15197 2015. | ISO 13485, Other | https://www.roche.com/
SD Biosensor, INC. | Gluconavii PRO BT | Blood glucose | Electrochemical | Amperometric | glucose dehydrogenase - flavin adenin dinucleotid (GDH - FAD) | Capillary blood, Venous blood, Arterial blood | fingertip and alternate site such as palm, upper arm, etc. | Adults, Children, Neonates | From medical trainee to lay user | 600mg/dL 33.3mmol/dL | 10mg/dL 0.6mmol/dL | <4% | No | 0.5 μL | 5 sec | 0-70% | -20 to 50°C | up to 93 | 2 to 32°C | 5.6 x 10.8 x 1.8 | 40 | 8°C | 45°C | LCD | 1000 | USB , Bluetooth 4.0 | 2x CR 2032 (3.0V coin cell battery) | 1000 | Yes | Yes | CE-IVDD | ISO 15197:2015 | ISO 13485 | https://www.sdbiosensor .com/
SD Biosensor, INC. | Gluconavii Pro | Blood glucose | Electrochemical | Amperometric | glucose dehydrogenase - flavin adenin dinucleotid (GDH - FAD) | Capillary blood, Venous blood, Arterial blood | fingertip and alternate site such as palm, upper arm, etc. | Adults, Children, Neonates | From medical trainee to lay user | 600mg/dL 33.3mmol/dL | 10mg/dL 0.6mmol/dL | <4% | No | 0.5 μL | 5 sec | 0-70% | -20 to 50°C | up to 93 | 2 to 32°C | 5.6 x 10.8 x 1.8 | 40 | 8°C | 45°C | LCD | 1000 | USB | 2x CR 2032 (3.0V coin cell battery) | 1000 | Yes | Yes | CE-IVDD | ISO 15197:2015 | ISO 13485 | https://www.sdbiosensor .com/
SD Biosensor, INC. | Codefree Plus BT | Blood glucose | Electrochemical | Amperometric | glucose oxidase (GOD) | Capillary whole blood | fingertip and alternate site such as palm, upper arm, etc. | Adults, Children | From medical trainee to lay user | 600mg/dL 33.3mmol/dL | 10mg/dL 0.6mmol/dL | <4% | No | 0.9 μL | 5 sec | 20-60% | -20 to 50°C | up to 93 | 2 to 32°C | 4.8 x 8.37 x 1.53 | 29.5 (without battery) | 10°C | 45°C | LCD | 300 | Bluetooth 4.0 | 1x CR 2032 (3.0V coin cell battery) | 1000 | No | Yes | CE-IVDD | ISO 15197:2015 | ISO 13485 | https://www.sdbiosensor .com/
SD Biosensor, INC. | Codefree Plus | Blood glucose | Electrochemical | Amperometric | glucose oxidase (GOD) | Capillary whole blood | fingertip and alternate site such as palm, upper arm, etc. | Adults, Children | From medical trainee to lay user | 600mg/dL 33.3mmol/dL | 10mg/dL 0.6mmol/dL | <4% | No | 0.9 μL | 5 sec | 20-60% | -20 to 50°C | up to 93 | 2 to 32°C | 4.8 x 8.37 x 1.53 | 29.5 (without battery) | 10°C | 45°C | LCD | 500 | N/A | 1x CR 2032 (3.0V coin cell battery) | 1000 | No | * | CE-IVDD | ISO 15197:2015 | ISO 13485 | https://www.sdbiosensor .com/
SD Biosensor, INC. | Gluconavii NFC | Blood glucose | Electrochemical | Amperometric | glucose dehydrogenase - flavin adenin dinucleotid (GDH - FAD) | Capillary blood, Venous blood, Arterial blood | fingertip and alternate site such as palm, upper arm, etc. | Adults, Children, Neonates | From medical trainee to lay user | 600mg/dL 33.3mmol/dL | 10mg/dL 0.6mmol/dL | <5% | No | 0.5 μL | 5 sec | 0-70% | -20 to 50°C | up to 93 | 2 to 32°C | 4.8 x 9.0 x 1.5 | 50 | 8°C | 45°C | LCD | 300 | USB, NFC | 1x CR 2032 (3.0V coin cell battery) | 1000 | No | Yes | CE-IVDD | ISO 15197:2015 | ISO 13485 | https://www.sdbiosensor .com/
SD Biosensor, INC. | Gluconavii GDH | Blood glucose | Electrochemical | Amperometric | glucose dehydrogenase - flavin adenin dinucleotid (GDH - FAD) | Capillary blood, Venous blood, Arterial blood | fingertip and alternate site such as palm, upper arm, etc. | Adults, Children, Neonates | From medical trainee to lay user | 600mg/dL 33.3mmol/dL | 10mg/dL 0.6mmol/dL | <5% | No | 0.5 μL | 5 sec | 0-70% | -20 to 50°C | up to 93 | 2 to 32°C | 4.8 x 9.0 x 1.5 | 50 | 8°C | 45°C | LCD | 500 | USB | 1x CR 2032 (3.0V coin cell battery) | 1000 | No | Yes | CE-IVDD | ISO 15197:2015 | ISO 13485 | https://www.sdbiosensor .com/
SD Biosensor, INC. | Mentor BT | Blood glucose | Electrochemical | Amperometric | glucose oxidase (GOD) | Capillary whole blood | fingertip and alternate site such as palm, upper arm, etc. | Adults, Children | From medical trainee to lay user | 600mg/dL 33.3mmol/dL | 10mg/dL 0.6mmol/dL | <4% | No | 0.3 μL | 5 sec | 20-60% | -20 to 50°C | up to 93 | 2 to 32°C | 5.0 x 9.32 x 1.82 | 50 | 10°C | 45°C | LCD | 500 | USB , Bluetooth 4.0 | 1x CR 2032 (3.0V coin cell battery) | 1000 | No | Yes | CE-IVDD, FDA 510k | ISO 15197:2015 | ISO 13485 | https://www.sdbiosensor .com/
SD Biosensor, INC. | Mentor | Blood glucose | Electrochemical | Amperometric | glucose oxidase (GOD) | Capillary whole blood | fingertip and alternate site such as palm, upper arm, etc. | Adults, Children | From medical trainee to lay user | 600mg/dL 33.3mmol/dL | 10mg/dL 0.6mmol/dL | <4% | No | 0.3 μL | 6 sec | 20-60% | -20 to 50°C | up to 93 | 2 to 32°C | 5.0 x 9.32 x 1.82 | 50 | 10°C | 45°C | LCD | 500 | USB | 1x CR 2032 (3.0V coin cell battery) | 1000 | No | Yes | CE-IVDD, FDA 510k | ISO 15197:2015 | ISO 13485 | https://www.sdbiosensor .com/
SD Biosensor, INC. | SD Codefree | Blood glucose | Electrochemical | Amperometric | glucose oxidase (GOD) | Capillary whole blood | fingertip and alternate site such as palm, upper arm, etc. | Adults, Children | From medical trainee to lay user | 600mg/dL 33.3mmol/dL | 10mg/dL 0.6mmol/dL | <4% | No | 0.9 μL | 7 sec | 20-60% | -20 to 50°C | up to 93 | 2 to 32°C | 4.7 x 95 x 1.75 | 47.5 (With Battery) | 10°C | 45°C | LCD | 500 | USB | 1x CR 2032 (3.0V coin cell battery) | 1000 | No | Yes | CE-IVDD | ISO 15197:2015 | ISO 13485 | https://www.sdbiosensor .com/
SD Biosensor, INC. | SD Check Gold 2 | Blood glucose | Electrochemical | Amperometric | glucose oxidase (GOD) | Capillary whole blood | fingertip and alternate site such as palm, upper arm, etc. | Adults, Children | From medical trainee to lay user | 600mg/dL 33.3mmol/dL | 10mg/dL 0.6mmol/dL | <4% | No | 0.9 μL | 8 sec | 20-60% | -20 to 60°C | up to 93 | 2 to 32°C | 4.44 x 9.2 x 1.9 | 50 | 10°C | 45°C | LCD | 400 | N/A | 1x CR 2032 (3.0V coin cell battery) | 1000 | No | * | Unknown | ISO 15197:2015 | ISO 13485 | https://www.sdbiosensor .com/
SD Biosensor, INC. | SD Check Gold | Blood glucose | Electrochemical | Amperometric | glucose oxidase (GOD) | Capillary whole blood | fingertip and alternate site such as palm, upper arm, etc. | Adults, Children | From medical trainee to lay user | 600mg/dL 33.3mmol/dL | 10mg/dL 0.6mmol/dL | <4% | No | 0.9 μL | 9 sec | 20-60% | -20 to 60°C | up to 93 | 2 to 32°C | 4.44 x 9.2 x 1.9 | 50 | 10°C | 45°C | LCD | 400 | N/A | 1x CR 2032 (3.0V coin cell battery) | 1000 | No | * | CE-IVDD, FDA 510k | ISO 15197:2015 | ISO 13485 | https://www.sdbiosensor .com/
TAIDOC TECHNOLOGY CORPORATION | td 4141a | Blood glucose, uric acid, total cholesterol | Electrochemical | Amperometric | glucose dehydrogenase - flavin adenin dinucleotid (GDH - FAD) | * | * | * | * | 800 mg/dL | 10 mg/dL | <5% | No | 0.5 μL | 5 sec | 0-70% | -20 to 60°C | * | 2 to 30°C | 9.78 x 3.35 x 1.8 | 30 | 10°C | 40°C | LCD | 1000 | Bluetooth, USB | 1x AAA batteries (1.5V) | * | * | Yes | Unknown | ISO 15197:2013 | ISO 13485, ISO 9001, MDSAP, FDA QSR, Other | https://taidoc.com/
TAIDOC TECHNOLOGY CORPORATION | td 4216 | β-Ketones Blood glucose, total cholesterol, triglycerides, uric acid, Hemoglobin, Lactate | Electrochemical | Amperometric | glucose dehydrogenase - flavin adenin dinucleotid (GDH - FAD) | * | * | * | * | 800 mg/dL | 10 mg/dL | <5% | * | 0.5 μL | 5 sec | 0-70% | -20 to 60°C | up to 90 | 2 to 30°C | 10.25 x 5.69 x 2.18 | 57 | 8°C | 45°C | LCD | 1000 | Bluetooth | 2x AAA batteries (1.5V) | * | * | Yes | Unknown | * | ISO 13485, ISO 9001, MDSAP, FDA QSR, Other | https://taidoc.com/
TAIDOC TECHNOLOGY CORPORATION | td 4289 | β-Ketones Blood glucose, Ketones, Cholesterol, Lactate, Uric acid | Electrochemical | Amperometric | glucose dehydrogenase - flavin adenin dinucleotid (GDH - FAD) | * | * | * | * | 800 mg/dL | 10 mg/dL | <5% | No | 0.5 μL | 5 sec | 0-70% | -20 to 60°C | up to 90 | 2 to 30°C | 9.63 x 6.1 x 2.2 | 58.3 | 8°C | 45°C | LCD | 1000 | Bluetooth, USB | 2x AAA batteries (1.5V) | * | * | Yes | Unknown | ISO 15197:2013 | ISO 13485, ISO 9001, MDSAP, FDA QSR, Other | https://taidoc.com/
TAIDOC TECHNOLOGY CORPORATION | td 4125 | Blood glucose | Electrochemical | * | glucose oxidase (GOD) | * | * | * | * | 600 mg/dL | 20 mg/dL | <5% | * | 0.7 μL | 7 sec | 30-55% | -20 to 60°C | up to 85 | 4 to 40°C | 9.7 x 5.2 x 1.6 | 47.8 | 10°C | 40°C | LCD | 450 | * | 2x AAA batteries (1.5V) | * | Yes | * | Unknown | * | ISO 13485, ISO 9001, MDSAP, FDA QSR, Other | https://taidoc.com/
TAIDOC TECHNOLOGY CORPORATION | td 4143 4140 | Blood glucose, β-Ketones , total cholesterol | * | * | * | * | * | * | * | 700 mg/dL | 10 mg/dL | <5% | No | 0.8 μL | 5 sec | 0-70% | -20 to 60°C | * | * | 5.5 x 4.03 x 1.04 | 11.8 | 10°C | 40°C | * | * | * | * | * | * | Yes | FDA 510k | * | ISO 13485, ISO 9001, MDSAP, FDA QSR, Other | https://taidoc.com/
TAIDOC TECHNOLOGY CORPORATION | td 4183 | Blood glucose | Electrochemical | Amperometric | glucose dehydrogenase - flavin adenin dinucleotid (GDH - FAD) | * | * | * | * | 650 mg/dL | 20 mg/dL | <5% | * | 0.5 μL | 6 sec | 20-65% | -20 to 60°C | up to 85 | 2 to 30°C | 9.03 x 5.23 x 1.8 | 46.4 | 10°C | 40°C | LCD | 400 | Bluetooth, USB | 1x AAA batteries (1.5V) | * | Yes | Yes | FDA 510k | ISO 15197:2013 | ISO 13485, ISO 9001, MDSAP, FDA QSR, Other | https://taidoc.com/
TAIDOC TECHNOLOGY CORPORATION | 4206A/4206B | Blood glucose, β-Ketones , total cholesterol | Electrochemical | Amperometric | glucose dehydrogenase - flavin adenin dinucleotid (GDH - FAD) | * | * | * | * | 700 mg/dL | 10 mg/dL | <5% | * | 0.8 μL | 5 sec | 0-70% | -20 to 60°C | up to 85 | 2 to 30°C | 9.7 x 6.22 x 2.89 | 57.8 | 10°C | 40°C | LCD | 1000 | A-strip port, B-Bluetooth | 2x AAA batteries (1.5V) | * | Yes | Yes | Unknown | * | ISO 13485, ISO 9001, MDSAP, FDA QSR, Other | https://taidoc.com/
TAIDOC TECHNOLOGY CORPORATION | td - 4235 | Blood glucose | Electrochemical | * | glucose oxidase (GOD) | * | * | * | * | 600 mg/dL | 20 mg/dL | <5% | No | 0.7 μL | 7 sec | 30-55% | -20 to 60°C | up to 85 | * | 9.3 x 3.6 x 2.13 | 45.4 | 10°C | 40°C | LCD | 150 | USB | 1x AAA batteries (1.5V) | * | * | Yes | Unknown | ISO 15197:2013 | ISO 13485, ISO 9001, MDSAP, FDA QSR, Other | https://taidoc.com/
TAIDOC TECHNOLOGY CORPORATION | td - 4257 | Blood glucose | Electrochemical | Amperometric | glucose dehydrogenase - flavin adenin dinucleotid (GDH - FAD) | * | * | * | * | 650 mg/dL | 20 mg/dL | <5% | * | 0.5 μL | 5 sec | 20-65% | -20 to 60°C | up to 85 | 2 to 30°C | 9.7 x 5.37 x 2.17 | 81 | 10°C | 40°C | * | 400 | * | 2x AAA batteries (1.5V) | * | Yes | * | FDA 510k | * | ISO 13485, ISO 9001, MDSAP, FDA QSR, Other | https://taidoc.com/
TAIDOC TECHNOLOGY CORPORATION | td - 4265 | Blood glucose | Electrochemical | * | glucose oxidase (GOD) | * | * | * | * | 600 mg/dL | 20 mg/dL | <5% | * | 0.5 μL | 5 sec | 30-50% | -20 to 60°C | * | 4 to 40°C | 12.5 x 3.3 x 1.4 | 39.8 | 10°C | 40°C | LCD | 150 | USB | 1x CR 2032 (3.0V coin cell battery) | * | Yes | Yes | FDA 510k | * | ISO 13485, ISO 9001, MDSAP, FDA QSR, Other | https://taidoc.com/
TAIDOC TECHNOLOGY CORPORATION | td - 4267 | Blood glucose | Electrochemical | * | glucose oxidase (GOD) | * | * | * | * | 600 mg/dL | 20 mg/dL | <5% | * | 0.5 μL | 5 sec | 30-55% | -20 to 60°C | * | 4 to 40°C | 9.0 x 5.5 x 1.8 | 49 | 10°C | 40°C | LCD | 450 | USB | 1x CR 2032 (3.0V coin cell battery) | * | Yes | Yes | FDA 510k | * | ISO 13485, ISO 9001, MDSAP, FDA QSR, Other | https://taidoc.com/
TAIDOC TECHNOLOGY CORPORATION | td - 4277 | Blood glucose | Electrochemical | Amperometric | glucose dehydrogenase - flavin adenin dinucleotid (GDH - FAD) | * | * | * | * | 600 mg/dL | 20 mg/dL | <5% | * | 0.5 μL | 5 sec | 35-60% | -20 to 60°C | up to 85 | 2 to 32°C | 9.6 x 6.1 x 2.6 | 67.2 | 10°C | 40°C | LCD | 1000 | USB, Bluetooth | 2x AAA batteries (1.5V) | * | Yes | Yes | FDA 510k | ISO 15197:2013 | ISO 13485, ISO 9001, MDSAP, FDA QSR, Other | https://taidoc.com/
TAIDOC TECHNOLOGY CORPORATION | td - 4277 advance | Blood glucose | Electrochemical | Amperometric | glucose dehydrogenase - flavin adenin dinucleotid (GDH - FAD) | * | * | * | * | 600 mg/dL | 20 mg/dL | <5% | * | 0.5 μL | 5 sec | 20-65% | -20 to 60°C | up to 85 | 2 to 32°C | 9.6 x 6.1 x 2.6 | 67.2 | 10°C | 40°C | LCD | 1000 | USB | 2x AAA batteries (1.5V) | * | Yes | Yes | Unknown | * | ISO 13485, ISO 9001, MDSAP, FDA QSR, Other | https://taidoc.com/
TAIDOC TECHNOLOGY CORPORATION | td - 4279c/4279d | Blood glucose | Electrochemical | Amperometric | glucose dehydrogenase - flavin adenin dinucleotid (GDH - FAD) | * | * | adults, children, neonates | * | 700 mg/dL | 10 mg/dL | <5% | * | 1.0 μL | 5 sec | 0-70% | -20 to 60°C | up to 85 | 2 to 30°C | 9.6 x 6.1 x 2.6 | 67.2 | 10°C | 40°C | LCD | 1000 | C-strip port, D-bluetooth | 2x AAA batteries (1.5V) | * | Yes | Yes | FDA 510k | * | ISO 13485, ISO 9001, MDSAP, FDA QSR, Other | https://taidoc.com/
TAIDOC TECHNOLOGY CORPORATION | td - 4280 | Blood glucose | Electrochemical | Amperometric | glucose dehydrogenase - flavin adenin dinucleotid (GDH - FAD) | * | * | * | * | 600 mg/dL | 20 mg/dL | <5% | * | 0.5 μL | 5 sec | 35-60% | -20 to 60°C | * | 2 to 32°C | 9.5 x 5.2 x 2.0 | 48.8 | 10°C | 40°C | LCD | 450 | Company provided cable | 2x AAA batteries (1.5V) | * | Yes | Yes | FDA 510k | * | ISO 13485, ISO 9001, MDSAP, FDA QSR, Other | https://taidoc.com/
TAIDOC TECHNOLOGY CORPORATION | td - 4283 | Blood glucose | Electrochemical | Amperometric | glucose dehydrogenase - flavin adenin dinucleotid (GDH - FAD) | * | * | * | * | 600 mg/dL | 20 mg/dL | <5% | * | 0.5 μL | 5 sec | 35-60% | -20 to 60°C | up to 85 | 2 to 32°C | 4.35 x 3.93 x 2.42 | 13 | 10°C | 40°C | LCD | 150 | Bluetooth, USB | 1x CR 2032 (3.0V coin cell battery) | * | * | Yes | Unknown | * | * | https://taidoc.com/
TAIDOC TECHNOLOGY CORPORATION | td - 4285 | Blood glucose | Electrochemical | Amperometric | glucose dehydrogenase - flavin adenin dinucleotid (GDH - FAD) | * | * | * | * | 650 mg/dL | 20 mg/dL | <5% | * | 0.5 μL | 6 sec | 20-65% | -20 to 60°C | * | 2 to 30°C | 9.6 x 6.0 x 1.5 | 41.2 | 10°C | 40°C | LCD | 450 | Company provided cable | 1x CR 2032 (3.0V coin cell battery) | * | Yes | Yes | Unknown | * | ISO 13485, ISO 9001, MDSAP, FDA QSR, Other | https://taidoc.com/
TAIDOC TECHNOLOGY CORPORATION | td - 4287 | Blood glucose | Electrochemical | Amperometric | glucose dehydrogenase - flavin adenin dinucleotid (GDH - FAD) | * | * | * | * | 600 mg/dL | 20 mg/dL | <5% | * | 0.5 μL | 5 sec | 35-60% | -20 to 60°C | * | 2 to 32°C | 10.34 x 2.81 x 1.51 | 31.8 | 10°C | 40°C | LCD | 150 | USB | 1x CR 2032 (3.0V coin cell battery) | * | * | * | Unknown | * | ISO 13485, ISO 9001, MDSAP, FDA QSR, Other | https://taidoc.com/
VivaChek Biotech (Hangzhou) Co., Ltd. | vivachek bravo | Blood glucose | * | * | glucose dehydrogenase (GDH) | capillary blood, venous blood | * | * | * | 600 mg/dL | 10 mg/dL | * | No | 0.6 μL | 5 sec | * | * | up to 90 | * | * | * | 5°C | 45°C | * | 500 | * | 1x CR 2032 (3.0V coin cell battery) | 1000 | Yes | * | Unknown | ISO 15197 2015. | * | https://www.vivachek.com/en/index.html
VivaChek Biotech (Hangzhou) Co., Ltd. | vivachek link plus | Blood glucose, β-Ketones , uric acid | Electrochemical | Biosensor | glucose oxidase (GOD), glucose dehydrogenase - flavin adenin dinucleotid (GDH - FAD) | capillary blood, arterial blood, venous blood | neonatal heel | adults, children, neonates | * | 600 mg/dL | 10 mg/dL | * | No | 0.5 μL | 5 sec | 0-70% | * | * | 2 to 30°C | * | * | 5°C | 45°C | * | * | * | * | * | * | Yes | Unknown | * | * | https://www.vivachek.com/en/index.html
VivaChek Biotech (Hangzhou) Co., Ltd. | vivachek plus | Blood glucose, β-Ketones , uric acid | Electrochemical | Biosensor | glucose oxidase (GOD), glucose dehydrogenase - flavin adenin dinucleotid (GDH - FAD) | capillary blood, arterial blood, venous blood | neonatal heel | adults, children, neonates | * | 600 mg/dL | 10 mg/dL | * | No | 0.5 μL | 5 sec | 0-70% | * | * | 2 to 30°C | * | * | 5°C | 45°C | * | 1000 | Bluetooth | * | * | Yes | Yes | Unknown | * | * | https://www.vivachek.com/en/index.html
VivaChek Biotech (Hangzhou) Co., Ltd. | fad plus | Blood glucose | Electrochemical | * | glucose dehydrogenase (GDH) | capillary blood, venous blood, arterial blood | * | * | * | 600 mg/dL | 10 mg/dL | * | * | 0.5 μL | 5 sec | 0-70% | * | up to 90 | * | * | * | 5°C | 45°C | * | 1000 | Bluetooth, USB | 2x AAA batteries (1.5V) | 500-1000 | Yes | Yes | Unknown | * | * | https://www.vivachek.com/en/index.html
VivaChek Biotech (Hangzhou) Co., Ltd. | fad smart | Blood glucose | Electrochemical | * | glucose dehydrogenase - flavin adenin dinucleotid (GDH - FAD) | capillary blood, arterial blood, venous blood | neonatal heel | adults, children, neonates | * | 600 mg/dL | 10 mg/dL | * | * | 0.5 μL | 5 sec | 0-70% | * | up to 90 | * | * | * | 5°C | 45°C | * | 500 | Bluetooth, USB | Rechargeable battery with capacity 250mAh | 150 | * | Yes | Unknown | * | * | https://www.vivachek.com/en/index.html
VivaChek Biotech (Hangzhou) Co., Ltd. | ino sound | Blood glucose | Electrochemical | * | glucose oxidase (GOD) | * | * | adults, children, neonates | * | 600 mg/dL | 10 mg/dL | * | No | 0.5 μL | 5 sec | 20-70% | * | up to 90 | * | * | * | 5°C | 45°C | * | 1000 | Bluetooth | 2x AAA batteries (1.5V) | 500 - 1000 | Yes | * | FDA 510k | * | * | https://www.vivachek.com/en/index.html
VivaChek Biotech (Hangzhou) Co., Ltd. | ino plus | Blood glucose | Electrochemical | * | glucose oxidase (GOD) | capillary blood, whole blood | * | adults, children, neonates | * | 600 mg/dL | 10 mg/dL | * | * | 0.5 μL | 5 sec | 20-70% | * | up to 90 | * | * | * | 5°C | 45°C | * | 1000 | Bluetooth, USB | 2x AAA batteries (1.5V) | 500 - 1000 | Yes | Yes | FDA 510k | * | * | https://www.vivachek.com/en/index.html
VivaChek Biotech (Hangzhou) Co., Ltd. | ino smart | Blood glucose | Electrochemical | * | glucose oxidase (GOD) | capillary blood, whole blood | * | * | * | 600 mg/dL | 10 mg/dL | * | * | 0.5 μL | 5 sec | 20-70% | * | up to 90 | * | * | * | 5°C | 45°C | * | 500 | Bluetooth, USB | Rechargeable battery with capacity 250mAh | 150 | * | Yes | FDA 510k | * | * | https://www.vivachek.com/en/index.html
VivaChek Biotech (Hangzhou) Co., Ltd. | fad | Blood glucose | Electrochemical | Amperometric | glucose dehydrogenase - flavin adenin dinucleotid (GDH - FAD) | capillary blood, arterial blood, venous blood | neonatal heel | * | * | 600 mg/dL | 10 mg/dL | * | No | 0.5 μL | 5 sec | 0-70% | * | up to 90 | * | * | * | 5°C | 45°C | * | 900 | USB | 1x CR 2032 (3.0V coin cell battery) | 1000+ | Yes | Yes | Unknown | ISO 15197 2015. | * | https://www.vivachek.com/en/index.html
VivaChek Biotech (Hangzhou) Co., Ltd. | eco | Blood glucose | Electrochemical | * | glucose oxidase (GOD) | capillary blood, whole blood | * | * | * | * | * | * | No | 0.5 μL | 5 sec | * | * | up to 90 | * | * | 39 | 5°C | 45°C | * | 300 | * | 2x CR 2032 (3.0V coin cell battery) | 1000+ | * | * | Unknown | ISO 15197 2015. | * | https://www.vivachek.com/en/index.html
VivaChek Biotech (Hangzhou) Co., Ltd. | ino | Blood glucose | Electrochemical | * | glucose oxidase (GOD) | capillary blood, whole blood | * | adults, children, neonates | * | * | * | * | No | 0.5 μL | 5 sec | 20-70% | * | up to 90 | * | * | * | 5°C | 45°C | * | 900 | USB | 3x CR 2032 (3.0V coin cell battery) | 1000+ | * | Yes | FDA 510k | ISO 15197 2015. | * | https://www.vivachek.com/en/index.html
Sinocare | Safe AQ UG | Blood glucose, uric acid | Electrochemical | Amperometric | * | capillary blood, whole blood, venous blood | * | adults, children | Self-administered, Trained staff in health care facilities | 600 mg/dL | 20 mg/dL | <6% | Yes | 0.6 μL | 6 sec | * | -20 to 55°C | * | * | 10.8 x 6.6 x 2.2 | 100 | 10°C | 35°C | LCD | 800 | Not applicable | 2x AAA alkaline batteries, 3V DC, 10mA | 1000 | Yes | * | CE-IVDD | * | ISO 13485, FDA QSR | https://en.sinocare.com/
Sinocare | Safe AQ Smart | Blood glucose | Electrochemical | Amperometric | * | capillary blood, whole blood, venous blood | fingertip | children, adults | Self-administered, Trained staff in health care facilities | 600 mg/dL | 20 mg/dL | <6% | * | 0.6 μL | 5 sec | * | -20 to 55°C | * | * | 8.4 x 6.0 x 2.6 | 73.5 | 10°C | 35°C | LCD | 10 | * | 2x AAA alkaline batteries, 3V DC, 50mA | * | Yes | * | CE-IVDD | * | ISO 13485, FDA QSR | https://en.sinocare.com/
Sinocare | Safe - Accu | Blood glucose | Electrochemical | Amperometric | * | capillary blood, whole blood, venous blood | fingertip | adults, children | Self-administered, Trained staff in health care facilities | 600 mg/dL | 20 mg/dL | <6% | * | 0.6 μL | 10 sec | * | -20 to 55°C | * | * | 9.5 x 5.5 x 1.8 | 40 | 10°C | 35°C | LCD | 200 | * | CR 2032 (3.0V coin cell battery), DC3V, 10mA | 1000 | Yes | * | CE-IVDD | * | ISO 13485 | https://en.sinocare.com/
Sinocare | Safe AQ Angel | Blood glucose | Electrochemical | Amperometric | * | capillary blood, whole blood, venous blood | fingertip | adults, children | Self-administered | 600 mg/dL | 20 mg/dL | <6% | No | 0.6 μL | 6 sec | * | -20 to 55°C | * | * | 10.3 x 5.7 x 2.2 | 52 | 10°C | 35°C | LCD | 200 | * | 2x AAA alkaline batteries, 3V DC | 1000 | Yes | * | CE-IVDD | ISO 15197:2013 | ISO 13485 | https://en.sinocare.com/
Sinocare | Safe - Accu 2 | Blood glucose | Electrochemical | Amperometric | * | capillary blood, whole blood, venous blood | fingertip | adults, children | Self-administered, Trained staff in health care facilities | * | * | * | * | 0.6 μL | 10 sec | * | -20 to 55°C | * | * | 10.5 x 6.0 x 2.3 | 77.2 | 10°C | 35°C | * | 200 | * | 2x AAA alkaline batteries, 3V DC, 10mA | * | Yes | * | CE-IVDD | * | ISO 13485 | https://en.sinocare.com/
Sinocare | Safe AQ Voice | Blood glucose | Electrochemical | Amperometric | * | capillary blood, whole blood, venous blood | fingertip | adults, children | Self-administered, Trained staff in health care facilities | 600 mg/dL | 20 mg/dL | <6% | * | 0.6 μL | 5 sec | * | -20 to 55°C | * | * | 8.4 x 6.0 x 2.6 | 73.5 | 10°C | 35°C | LCD | 10 | * | 2x AAA alkaline batteries, 3V DC, 50mA | * | Yes | * | CE-IVDD | * | ISO 13485 | https://en.sinocare.com/
Sinocare | Safe AQ Air | Blood glucose | Electrochemical | Amperometric | * | capillary blood, whole blood, venous blood | fingertip | adults, children | Self-administered, Trained staff in health care facilities | 600 mg/dL | 20 mg/dL | <6% | * | 0.6 μL | 5 sec | * | -20 to 55°C | * | * | 8.4 x 6.0 x 2.6 | 73.5 | 10°C | 35°C | LCD | 10 | Bluetooth | 2x AAA alkaline batteries, 3V DC, 50mA | * | Yes | Yes | Unknown | * | ISO 13485 | https://en.sinocare.com/
Sinocare | Safe AQ Pro I | Blood glucose | Electrochemical | Amperometric | glucose dehydrogenase - flavin adenin dinucleotid (GDH - FAD) | capillary blood, whole blood, venous blood | fingertip | adults, children | Self-administered, Trained staff in health care facilities | 600 mg/dL | 20 mg/dL | <2.8% | No | 1.0 μL | 5 sec | 10-70% | -20 to 55°C | * | 5 to 35°C, <95% RH | 8.5 x 5.3 x 2.36 | 53.8 | 10°C | 40°C | LCD | 500 | * | 2x AAA batteries (1.5V), #V DC, 20mA | * | Yes | * | CE-IVDD | ISO 15197:2013 | ISO 13485 | https://en.sinocare.com/
A. Menarini Industrie Farmaceutiche Riunite S.r.l. | GlucoMen areo | Blood glucose | Electrochemical | Amperometric | glucose oxidase (GOD) | capillary blood, whole blood | palm, forearm, fingertip | adults, children | Self-administered | 33.3 mmol/L | 1.1 mmol/L | * | No | 0.5 μL | 5 sec | 10-70% | -20 to 50°C | up to 90 | 4 to 30°C, 20-90% RH | 8.55 x 5.6 x 1.82 | 46 | 5°C | 45°C | LCD | 730 | NFC, Bluetooth | 2x CR 2032 (3.0V coin cell battery) | 1000 | No | Yes | CE-IVDD | ISO 15197 2015. | * | https://glucomenday.com/newplatform/
A. Menarini Industrie Farmaceutiche Riunite S.r.l. | Glucofix premium | Blood glucose, β-Ketones | Electrochemical | Biosensor | glucose oxidase (GOD) | * | * | * | * | 600 mg/dL | 20 mg/dL | * | * | 0.3 μL | 4 sec | 25-60% | * | up to 90 | * | 5.8 x 9.8 x 1.7 | 75 | 5°C | 45°C | LCD | 400 | * | * | * | * | * | CE-IVDD | * | * | https://glucomenday.com/newplatform/
A. Menarini Industrie Farmaceutiche Riunite S.r.l. | Glucomen LX2 | Blood glucose, β-Ketones | Electrochemical | Biosensor | glucose oxidase (GOD) | capillary blood, whole blood | fingertip, palm, forearm | adults, children | Self-administered, Trained staff in health care facilities | 33.3 mmol/L | 1.1 mmol/L | * | No | 0.3 μL | 4 sec | 25-60% | -25 to 46°C | up to 90 | 4 to 30°C, 10-90% RH | 5.8 x 9.8 x 1.7 | 75 | 5°C | 45°C | LCD | 400 | Company provided cable | 2x CR 2032 (3.0V coin cell battery) | 1000 | Yes | Yes | CE-IVDD | * | * | https://glucomenday.com/newplatform/
A. Menarini Industrie Farmaceutiche Riunite S.r.l. | Glucomen LX plus | Blood glucose, β-Ketones | Electrochemical | Biosensor | glucose oxidase (GOD) | capillary blood, whole blood | fingertip, palm, forearm | adults, children | Self-administered, Trained staff in health care facilities | 33.3 mmol/L | 1.1 mmol/L | * | No | 0.3 μL | 4 sec | 25-60% | -25 to 46°C | up to 90 | 4 to 30°C, 10-90% RH | 5.8 x 9.8 x 1.7 | 75 | 5°C | 45°C | LCD | 400 | Company provided cable | * | * | Yes | Yes | CE-IVDD | * | * | https://glucomenday.com/newplatform/
A. Menarini Industrie Farmaceutiche Riunite S.r.l. | GlucoMen areo 2k | Blood glucose, β-Ketones | Electrochemical | Amperometric | glucose oxidase (GOD) | capillary blood, whole blood | fingertip, palm, forearm | adults, children | Self-administered, Trained staff in health care facilities | 33.3 mmol/L | 1.1 mmol/L | * | No | 0.5 μL | 5 sec | 10-70% | -20 to 50°C | up to 90 | 4 to 30°C | 8.55 x 5.6 x 1.82 | 46 | 5°C | 45°C | LCD | 730 | NFC, Bluetooth | 2x CR 2032 (3.0V coin cell battery) | 900 | Yes | Yes | CE-IVDD | * | * | https://glucomenday.com/newplatform/
A. Menarini Industrie Farmaceutiche Riunite S.r.l. | GlucoFix tech | Blood glucose | Electrochemical | Amperometric | glucose oxidase (GOD) | capillary blood, whole blood | fingertip, palm, forearm | adults, children | Self-administered, Trained staff in health care facilities | 600 mg/dL | 20 mg/dL | * | No | 0.5 μL | 5 sec | 10-70% | -20 to 50°C | up to 90 | 4 to 30°C, 20-90% RH | 8.55 x 5.6 x 1.82 | 46 | 5°C | 45°C | LCD | 730 | NFC, Bluetooth | 2x CR 2032 (3.0V coin cell battery) | 1000 | Yes | Yes | CE-IVDD | ISO 15197 2015. | * | https://glucomenday.com/newplatform/
A. Menarini Industrie Farmaceutiche Riunite S.r.l. | GlucoFic tech 2k | Blood glucose, β-Ketones | Electrochemical | Amperometric | glucose oxidase (GOD) | capillary blood, whole blood | fingertip, palm, forearm | adults, children | Self-administered, Trained staff in health care facilities | 600 mg/dL | 20 mg/dL | * | No | 0.5 μL | 5 sec | 10-70% | -20 to 50°C | up to 90 | 4 to 30°C, 20-90% RH | 8.55 x 5.6 x 1.82 | 46 | 5°C | 45°C | LCD | 730 | NFC, Bluetooth | 2x CR 2032 (3.0V coin cell battery) | 900 | Yes | Yes | CE-IVDD | ISO 15197 2015. | * | https://glucomenday.com/newplatform/
A. Menarini Industrie Farmaceutiche Riunite S.r.l. | Glucomen Ready | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | up to 90 | * | 10.7 x 6.0 x 2.0 | 95 | * | * | * | * | * | * | * | * | * | CE-IVDD | * | * | https://glucomenday.com/newplatform/
Note on W127: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: @Lejla Druskic please have a look here Reply: done