## Sheet1
Manufacturer | Model name | Intended use | Adjunct or non-adjunct | Calibration requirements | Real-time or inttermittently scanned | Age limit | Sensor duration | Sensor Placement site | Sensor application procedure | Sensor/transmitter configuration | Transmitter life time | Transmitter power source | Sensor&transmitter operating temperature | External reader/reading device | Mobile phone application (Android/iOS) | Type of connectivity | Sharing features available with 3rd party | Measurement range | Detection range unit | MARD in % | Alarms | Software solution/platform to consolidate results for healtcare provider and patient | Regulatory approval | Standards for the manufacturer | Standards for product performance | Organization main website
Abbott | Libre 1 | - measurement of interstitial fluid glucose levels in people with diabetes mellitus (including pregnant women) - designed to replace blood glucose testing in the self-management of diabetes, including dosing of insulin | non - adjunct | not required | real-time | real-time | 14 days | back of upper arm | single use applicator | no transmitter | * | * | 10-45°C | yes | yes | NFC | yes | 1-27.8mmol/L (20-500mg/dL) | adjustable | < 10.0 | no | yes | CE-IVDD | ISO 13485:2016 | * | https://www.freestyle.abbott/us-en/home.html
Abbott | Libre 2 | - measurement of interstitial fluid glucose levels in people with diabetes mellitus (including pregnant women) - designed to replace blood glucose testing in the self-management of diabetes, including dosing of insulin - detection of trends and tracking of patterns to aid in the detection of episodes of hyperglycemia and hypoglycemia | non - adjunct | not required | real-time | 4+ | 14 days | back of upper arm | single use applicator | integrated transmitter | duration of sensor | integrated battery | 10-45°C | yes | yes | NFC, bluetooth *only for alarms | yes | 1-27.8mmol/L (20-500mg/dL) | adjustable | 8.2 | yes | yes | CE-IVDD, FDA 510k, Brazil ANVISA, Australia ARTG | ISO 13485:2016 | * | https://www.freestyle.abbott/us-en/home.html
Abbott | Libre 3 | - measurement of interstitial fluid glucose levels in people with diabetes mellitus (including pregnant women) - designed to replace blood glucose testing in the self-management of diabetes, including dosing of insulin - detection of trends and tracking of patterns to aid in the detection of episodes of hyperglycemia and hypoglycemia | non - adjunct | not required | real-time | 4+ | 14 days | back of upper arm | single use applicator | integrated transmitter | duration of sensor | integrated battery | 10-45°C | yes | yes | bluetooth | yes | 2.2-22.2mmol/L (40-400mg/dL) | adjustable | 7.9 | yes | yes | CE-IVDD, FDA-510k | ISO 13485:2016 | ISO 15197:2013 | https://www.freestyle.abbott/us-en/home.html
Abbott | Libre pro | - indicated for detecting trends and tracking patterns and glucose level excursions above or below the desired range, facilitating therapy adjustments in persons with diabetes - intended for use by health care professionals and requires a prescription. | adjunct | not required | no scanning/data is hidden | 18+ | 14 days | back of upper arm | single use applicator | no transmitter | * | * | * | yes | no | nfc | yes | * | adjustable | * | no | yes | CE-IVDD | ISO 13485:2016 | * | https://www.freestyle.abbott/us-en/home.html
Abbott | Lingo | - measurement of various health metrics in health-conscious individuals | * | * | real-time | * | 14 days | * | * | * | * | * | * | * | yes | * | yes | 3.0-11.1mmol/L (55-200mg/dL) | * | * | * | * | FDA-510k | ISO 13485:2016 | * | https://www.freestyle.abbott/us-en/home.html
Abbott | Freestyle Rio | - measurement of interstitial fluid glucose levels in people with diabetes mellitus | * | * | real-time | * | 14 days | * | * | * | * | * | * | * | yes | * | yes | 2.2-22.2mmol/L (40-400mg/dL) | * | * | * | * | * | ISO 13485:2016 | * | https://www.freestyle.abbott/us-en/home.html
Dexcom | G6 | - measurement of interstitial fluid glucose levels in people with diabetes mellitus - designed to replace blood glucose testing in the self-management of diabetes, including dosing of insulin - detection of trends and tracking of patterns to aid in the detection of episodes of hyperglycemia and hypoglycemia | non - adjunct | not required | real-time | 2+ | 10 days | back of upper arm, abdomen, upper buttocks | single use applicator | separate transmitter | 90 days | integrated battery | 10-45°C | yes | yes | bluetooth | yes | 2.2-22.2mmol/L (40-400mg/dL) | adjustable (depends on location) | 9 | yes | yes | CE-IVDD, FDA-510k | ISO 13485:2016 ISO 14971:2019 | ISO 15197:2013 | https://www.dexcom.com/
Dexcom | G7 | - measurement of interstitial fluid glucose levels in people with diabetes mellitus (including pregnant women) - designed to replace blood glucose testing in the self-management of diabetes, including dosing of insulin - detection of trends and tracking of patterns to aid in the detection of episodes of hyperglycemia and hypoglycemia | non - adjunct | not required | real-time | 2+ | 10 days + 12 hours grace period | back of upper arm, abdomen, upper buttocks | single use applicator | integrated transmitter | duration of sensor | integrated battery | 10-42°C | yes | yes | bluetooth | yes | 2.2-22.2mmol/L (40-400mg/dL) | adjustable (depends on location) | 8.2 | yes | yes | CE-IVDD, FDA-510k, Australia ARTG, Health Canada | ISO 13485:2016 ISO 14971:2019 | ISO 15197:2013 | https://www.dexcom.com/
Dexcom | ONE | - measurement of interstitial fluid glucose levels in people with diabetes mellitus (including pregnant women) - designed to replace blood glucose testing in the self-management of diabetes, including dosing of insulin - detection of trends and tracking of patterns to aid in the detection of episodes of hyperglycemia and hypoglycemia | non - adjunct | not required | real-time | 2+ | 10 days | back of upper arm, abdomen, upper buttocks | single use applicator | separate transmitter | 90 days | integrated battery | 10-45°C | yes | yes | bluetooth | yes | 2.2-22.2mmol/L (40-400mg/dL) | adjustable (depends on location) | 9 | yes | yes | CE-IVDD | ISO 13485:2016 ISO 14971:2019 | ISO 15197:2013 | https://www.dexcom.com/
Dexcom | ONE + | - measurement of interstitial fluid glucose levels in people with diabetes mellitus (including pregnant women) - designed to replace blood glucose testing in the self-management of diabetes, including dosing of insulin - detection of trends and tracking of patterns to aid in the detection of episodes of hyperglycemia and hypoglycemia | non - adjunct | not required | real-time | 2+ | 10 days + 12 hours grace period | back of upper arm, abdomen, upper buttocks (depends on age) | single use applicator | integrated transmitter | duration of sensor | integrated battery | 10-42°C | yes | yes | bluetooth | yes | 2.2-22.2mmol/L (40-400mg/dL) | adjustable (depends on location) | 9 | yes | yes | CE-IVDD | ISO 13485:2016 ISO 14971:2019 | ISO 15197:2013 | https://www.dexcom.com/
Dexcom | Stelo | - measurement of interstitial fluid glucose levels in people with diabetes mellitus | * | * | real-time | * | 10 days | * | * | * | * | * | * | * | yes | * | yes | 3.8-13.8mmol/L (70-250mg/dL) | * | * | * | yes | FDA-510k | ISO 13485:2016 ISO 14971:2019 | * | https://www.dexcom.com/
Microtech | GlucoRx AiDEX | - measurement of subcutaneous tissue glucose levels in people with diabetes mellitus - detection of trends and tracking of patterns to aid in the detection of episodes of hyperglycemia and hypoglycemia - reminder of when fingertip testing is required. | adjunct | optional | real-time | 14+ | 14 days | outside and back of upper arm, abdomen | single use applicator | separate transmitter | up to 4 years | sensor battery | 5-40°C | no | yes | bluetooth | yes | 2.0-25mmol/L (36-450mg/dL) | adjustable (depends on location) | 9.08 | no | yes | CE-IVDD | ISO 13485:2016 ISO 14971:2019 | ISO 15197:2013 | https://aidexcgm.com/
Medtrum | TouchCare Nano CGM | - measurement of interstitial fluid glucose levels in people with diabetes mellitus - designed to replace blood glucose testing in the self-management of diabetes - detection of trends and tracking of patterns to aid in the detection of episodes of hyperglycemia and hypoglycemia | non - adjunct | required | real-time | 2+ | 14 days | back of upper arm, abdomen, approximately 2.5 cm (1 inch) away from the insulin infusion site | single use applicator | separate transmitter | around 90 days | rechargable battery | 10-45°C | yes | yes | bluetooth | yes | 2.2-22.2mmol/L (40-400mg/dL) | adjustable (depends on location) | 9.2 | yes | yes | CE-IVDD | ISO 13485:2016 ISO 14971:2019 | ISO 15197:2013 | https://www.medtrum.com/
POCtech | CT10 | - measurement of interstitial fluid glucose levels in people with diabetes mellitus - detection of trends and tracking of patterns - used in combination with figertip pricking | adjunct | required | real-time | 14+ | 10 days | back of upper arm or middle abdomen | single use applicator | separate transmitter | 1 year | CR1620 battery | 5-40°C | yes | yes | bluetooth | yes | 1.7-25mmol/L (30-450 mg/dL); | adjustable (depends on location) | 9.61 | yes | yes | CE-IVDD | ISO 13485:2016 ISO 14971:2019 | ISO 15197:2013 | https://cgm-poctech.me/en/index.html
A.MENARINI diagnostics | GlucoMen Day CGM | - measurement of interstitial fluid glucose levels in people with diabetes mellitus - designed to replace blood glucose testing in the self-management of diabetes - detection of trends and tracking of patterns to aid in the detection of episodes of hyperglycemia and hypoglycemia | non - adjunct | required | real-time | 6+ | 14 days | back of upper arm, abdomen, upper buttocks | reusable tool | separate transmitter | 5 years | 1 non-serviceable, rechargeable lithium-ion polymer battery | 5-45°C | no | app | bluetooth | yes | 2.2-22.2mmol/L (40-400mg/dL) | adjustable | 9.7 | yes | yes | CE-IVDD | ISO 13485:2016 ISO 14971:2019 | ISO 15197:2013 | https://glucomenday.com/newplatform/glucomen-day-cgm/
SI Bionics | GS1 CGM | - measurement of interstitial fluid glucose levels in people with diabetes mellitus - designed to replace blood glucose testing in the self-management of diabetes - detection of trends and tracking of patterns to aid in the detection of episodes of hyperglycemia and hypoglycemia | non-adjunct | not required | real-time | 18+ | 14 days | back of upper arm | single use applicator | integrated transmitter | duration of sensor | sensor battery | 5-40°C | no | yes | bluetooth | yes | 2.2-25mmol/L (40-450mg/dL) | adjustable | 8.83 | yes | yes | CE-MDR | * | * | https://www.sibionicscgm.com/
i-SENS | CareSens Air | - measurement of interstitial fluid glucose levels in people with diabetes mellitus - detection of trends and tracking of patterns - used in combination with figertip pricking | adjunct | required | real-time | 18+ | 15 days | back of upper arm | single use applicator | integrated transmitter | duration of sensor | integrated battery | 10-42°C | no | yes | bluetooth | yes | 2.2-28mmol/L (40-500mg/dL) | adjustable | 12.2 | yes | yes | CE-IVDD | * | * | https://i-sens.com/
Sinocare | iCan CGM | - measurement of interstitial fluid glucose levels in people with diabetes mellitus - designed to replace blood glucose testing in the self-management of diabetes - detection of trends and tracking of patterns to aid in the detection of episodes of hyperglycemia and hypoglycemia | non - adjunct | not required | real-time | 18+ | 15 days | abdomen and other sites approved by HCP's | single use applicator | separate transmitter | duration of sensor | 1 non-serviceable, non-rechargeable button cell inside the transmitter, DC 1.5 V | 10-42°C | no | yes | bluetooth | yes | 2.0-25mmol/L (36-450mg/dL) | adjustable | 8.71 | yes | yes | CE-IVDD | * | * | https://www.icancgm.com/
Ascensia diabetes | Eversense E3 CGM | - measurement of interstitial fluid glucose levels in people with diabetes mellitus - designed to replace blood glucose testing in the self-management of diabetes - detection of trends and tracking of patterns to aid in the detection of episodes of hyperglycemia and hypoglycemia | non - adjunct | required | real-time | 18+ | 180 days | back of upper arm | inserted under skin through a medical procedure | separate transmitter | up to 1 year | rechargable battery | 5-40°C | no | yes | bluetooth | yes | 2.2-22.2mmol/L (40-400mg/dL) | adjustable | 8.5 | yes | yes | * | * | * | https://www.ascensiadiabetes.com/eversense/
Medtronic | Simplera | - measurement of interstitial fluid glucose levels in people with diabetes mellitus - designed to replace blood glucose testing in the self-management of diabetes - detection of trends and tracking of patterns to aid in the detection of episodes of hyperglycemia and hypoglycemia | non - adjunct | not required | real-time | 2+ | 7 days | back of upper arm, upper buttocks (ages 2-17) | single use applicator | integrated transmitter | duration of sensor | integrated battery | 2-40°C | no | yes | bluetooth | yes | x-22.1mmol/L (x-400mg/dL) | adjustable | 10.1 | yes | yes | CE-IVDD | ISO 13485:2016 | * | https://www.medtronic.com/us-en/index.html
Medtronic | Guardian 4 Connect CGM | - measurement of interstitial fluid glucose levels in people with diabetes mellitus - used in combination with figertip pricking | adjunct | not required | real-time | 7+ | 7 days | back of upper arm | reusable tool | separate transmitter | 7 days when fully charged | rechargable battery | * | no | yes | bluetooth | yes | * | adjustable | 10.7 | yes | yes | CE-IVDD | ISO 13485:2016 | ISO 15197:2013 | https://www.medtronic.com/us-en/index.html
Note on U2: Lejla Druskic: Bailey, T., Bode, B. W., Christiansen, M. P., Klaff, L. J., & Alva, S. (2015). The performance and usability of a factory-calibrated flash glucose monitoring system. Diabetes Technology & Therapeutics, 17(11), 787-794. DOI: 10.1089/dia.2015.0183 This study comprehensively evaluates the FreeStyle Libre 1 system and reports the MARD as approximately 11.4%.
Note on U3: Lejla Druskic: source https://www.support.freestyle.abbott/hc/en-us/articles/21906401259665-How-accurate-is-my-FreeStyle-Libre-2-Plus-sensor
Note on O14: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: @Lejla Druskic what does it mean? Reply: It's a system where you also have an insulin pump and there is a device that controls it. That device is called a PDM - personal diabetes manager and also serves as a reciever for the values. Reply: so can we write "yes (within the insulin pump controller)" Reply: yes
Note on U21: Lejla Druskic: https://www.barndiabetesfonden.se/for-drabbade/hjalpmedel/glukossensor/sensor-jamfor/
Note on U22: Lejla Druskic: https://www.ncbi.nlm.nih.gov/pmc/articles/PMC10675113/