## Sheet1
Company | Company HQ Country | WEB URL | Focal contact | Equipment image (main X-ray generator) | Equipment Name / Model | Portable Digital Radiography system, consisting of X-ray generator & flat panel detector, capable of taking high quality CXR in erect position (with dedicated generator & detector stand) | Total weight of entire system including carrying case | Equipment assembly time (set up time required for removal from tranportation case and make the equipment ready for use) | X-ray Generator | Output power | kV range & step size | mAs range | Maximum mA at 70 kV | Exposure time range | Anode type and focal spot size | Heat storage capacity | Overload protection (Y/N + details if Y) | Collimator (type, material, positioning light) | Total filtration at 70 kV and above | Generator mount / stand (Y, built into device Vs Y, separate component / N + details ) | X-ray source to detector distance for CXR in erect position (SID) | Generator to detector alignment with cranial/caudal beam angulation (Y/N + details if Y) | Weight (X-ray generator) | Input Power (Battery / AC) | Battery charging time (Not applicable for AC powered units) | Number of CXR exposures when used with fully charged battery (unlimited exposures for AC powered units) | In-use battery charging (Y/N) (Not applicable for AC powered units) | Extra battery (Y/N + details if Y) (Not applicable for AC powered units) | Low-battery indicator (Y/N) (Not applicable for AC powered units) | Exposure release switch (type and details) | Storage conditions (Temperature in degree celcius, Humidity in %, atmospheric pressure in hPa) | Use/Operating condition (Temperature in degree celcius, Humidity in %, atmospheric pressure in hPa) | Shock/drop test, water/dust penetration test | X-ray Detector | Active detector area | Time to display image after exposure | Pixel pitch ( µm) | Spatial resolution (lp/mm) | DQE, detection quantum effiency (%) | Dynamic range (bit) | Power source (Battery / AC) | Recharge power source (applicable for battery run detectors) | Exposure capacity (number of exposures when the detector is fully charged) | Charging time (applicable for battery run detectors) | in-use battery charging (Y/N) (applicable for battery run detectors) | Extra battery (Y/N + details if Y) | Connection type (wired / wireless) | Weight (detector only) | Detector software (for image processing) | Storage condition (Temperature in degree celcius, Humidity in %, atmospheric pressure in hPa) | Use/Operating condition (Temperature in degree celcius, Humidity in %, atmospheric pressure in hPa) | Shock/drop test, water/dust penetration test | Detector stand/frame for taking CXR in erect position | DR Workstation | Workstation device type (eg; laptop or tablet) and weight | Display | Display Languages | Processor | RAM | Hard drive | Transfer data to other workstations / PC-consoles/networks | Storage condition (Temperature in degree celcius, Humidity in %, atmospheric pressure in hPa) | Use/Operating condition (Temperature in degree celcius, Humidity in %, atmospheric pressure in hPa) | Shock/drop test, water/dust penetration test | Software | DICOM 3.0 | Patient registration and data (Y/N) | Exposure parameter regulation (Y/N + details if Y) | image post-processing (Y/N + details if Y) | APR Chest included (Y/N) | CAD/AI integration for CXR TB screening (Y/N + details if Y) | Transportation case (Y/N + details if Y) | Radiation Protection | Leakage radiation at 1 m and 2 m from source (in µSv) | Lead apron for operator (Y/N + details if Y) | Lead apron for patient (Y/N + details if Y) | Regulatory Requirements | Regulatory approval by specific agencies | ISO 13485 certification (medical devices) | IEC 60601 certification (medical electrical equipment) | Documentation | User manual and maintenance/service manual in English (includes safety recommendations) (Y/N) | List of important spare parts and accessories with their part number and costing (Y/N) | Certificate of calibration and inspection from factory (Y/N) | Initial warranty contract (Y/N + details if Y) | Logbook with instruction for daily, weekly, monthly and quarterly maintenance checklist (Y/N) | Terms of Reference for hospital service technician and company service engineer (Y/N) | Installation and Training | Application training to be carried out for 2 days for the radiographers and attendants using the machine installed at the hospital facility after commissioning the x-ray unit and biomedical technicians on basic maintenance and troubleshooting (Y/N) | Provisions for centralized training for multiple non-facility based installations procured at the same time (Y/N) | Service and maintenance | Comprehensive guarantee and service agreements (Y/N) | Experience with TB programs in LMICs or UMICs (Y/N) | Estimated cost (USD) per unit
1st SOURCE | United States | https://1stsourceimaging.com/iq-flex-m_portable-dr/ | sales@fsimed.com |  | IQ Flex-M | Yes | contact manufacturer | contact manufacturer |  | 2.5 kW | 40-100 kV (1 kV steps) | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | AC, 90 -220 V | Not applicable | unlimited exposures on AC | Not applicable | Not applicable | Not applicable | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer |  | 43 x 35 cm 43 x 43 cm | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer |  | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer |  | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer |  | contact manufacturer | contact manufacturer | contact manufacturer |  | contact manufacturer | contact manufacturer | contact manufacturer |  | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer |  | contact manufacturer | contact manufacturer |  | contact manufacturer | contact manufacturer | contact manufacturer
ALERIO | India | https://www.alerio.in/1-6kw-portable-x-ray-smart-1600/ | sales@alerio.in office@iatome.in |  | ALERIO Smart 1600 | Yes | 10 kg | 10 minutes |  | 1.6 kW | 40 - 90 kV (1 kV Steps) | 1 - 40 mAs | 23 mA | 0.01 - 2 sec | Stationary, 0.8 mm focal spot | 3.6 kHU | Yes; The unit is programmed with overload protection, an error message is generated during overload condition without the X-ray being exposed | Manual , built-in laser guided LED light | > 2.0 mmAl | Yes; separate component. Manual crank type stand, 4 wheels for ease of accessibility and 2 with brakes | 100 cm | contact manufacturer | 8.5 kg | Battery, 24 VDC | 3 hours | > 70 exposures | No | Yes, One additional battery pack | Yes | Wired (2 m) and wireless remote control hand switch | Temperature: 10 to 40 °C Humidity: 30 - 75% (non-condensing) Atmospheric pressure: 700 hPa - 1060 hPa | Temperature: -10 to 70°C Humidity: 5 - 95% (non-condensing) Atmospheric pressure: 500 - 1100 hPa | contact manufacturer |  | 43 x 35 cm (CareView® 1500Cwe) | 2 - 4 sec | 140 µm | 3.57 lp/mm | 71% at 0 lp/mm, 61% at 0.5 lp/mm 54% at 1 lp/mm, 46% at 1.5 lp/mm 40% at 2 lp/mm, 31% at 2.5 lp/mm 23% at 3 lp/mm, 16% at 3.5 lp/mm | contact manufacturer | Battery | contact manufacturer | >200 | 2.5 hours | Yes | Yes, One additional Lithium Polymer battery | wireless | 3.7 kg | Yes DROC | contact manufacturer | contact manufacturer | IP64 for dust and water-resistance | contact manufacturer |  | Laptop/tablet: <3 Kg | LCD | English | Intel i5 | 8 GB | 1 TB | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer |  | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer |  | < 100 µSv/hr Measuring conditions: at 1 meter from Focal Spot | contact manufacturer | contact manufacturer |  | contact manufacturer | Yes | Yes |  | Yes | Yes | Yes | Yes 2 Years warranty | Contact manufacturer | Contact manufacturer |  | Contact manufacturer | Contact manufacturer |  | Contact manufacturer | Contact manufacturer | contact manufacturer
ALERIO | India | https://www.alerio.in/ | sales@alerio.in office@iatome.in | contact manufacturer | ALERIO Smart 2800 | Yes | 12 kg approx. (X-ray generator only) | 10 minutes |  | 2.8 kW | 40 - 100 kV (1kV steps) | 1 - 100 mAs | 40 mA | 0.01 - 2 sec | Stationary, 1.2 mm focal spot | 5.8 kHU | Yes; The unit is programmed with overload protection, an error message is generated during overload condition without the X-ray being exposed | Manual , built-in laser guided LED light | > 2.0 mm Al | Yes; separate component. Manual crank type stand, 4 wheels for ease of accessibility and 2 with brakes | 100 cm | contact manufacturer | <12 kg | AC, 180 - 260 V | Not applicable | unlimited exposures on AC | Not applicable | Not applicable | Not applicable | Wired (2 m) and wireless remote control hand switch | Temperature: 10 to 40 °C Humidity: 30 - 75% (non-condensing) Atmospheric pressure: 700 hPa - 1060 hPa | Temperature: -10 to 70°C Humidity: 5 - 95% (non-condensing) Atmospheric pressure: 500 - 1100 hPa | contact manufacturer |  | 43 x 35 cm (CareView® 1500Cwe) | 2 - 4 sec | 140 µm | 3.57 lp/mm | 71% at 0 lp/mm, 61% at 0.5 lp/mm 54% at 1 lp/mm, 46% at 1.5 lp/mm 40% at 2 lp/mm, 31% at 2.5 lp/mm 23% at 3 lp/mm, 16% at 3.5 lp/mm | contact manufacturer | AC (100 -250V) | Not applicable | continuous exposure | Not applicable | Not applicable | Not applicable | wireless | 3.7 kg | Yes DROC | contact manufacturer | contact manufacturer | IP64 for dust and water-resistance | contact manufacturer |  | Laptop/tablet: <3 Kg | LCD | English | Intel i5 | 8 GB | 1 TB | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer |  | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer |  | < 100 µSv/hr Measuring conditions: at 1 meter from Focal Spot | contact manufacturer | contact manufacturer |  | contact manufacturer | Yes | contact manufacturer |  | Yes | Yes | Yes | Yes 2 Years warranty | Contact manufacturer | Contact manufacturer |  | Contact manufacturer | Contact manufacturer |  | Contact manufacturer | Contact manufacturer | contact manufacturer
ALERIO | India | https://www.alerio.in/4-2kw-mobile-x-ray-smart-4200/ | sales@alerio.in office@iatome.in |  | ALERIO Smart 4200 | Yes | 80 kg approx. (X-ray generator only) | Already assembled |  | 4.2 kW | 40 - 120 kV (1kV steps) | 1 - 250 mAs | 52 mA | 0.01 - 5 sec | Stationary, 1.8mm and 0.6 mm dual focal spot | 17.4 kHU | Yes; The unit is programmed with overload protection, an error message is generated during overload condition without the X-ray being exposed | Manual , built-in laser guided LED light | > 2.0 mm Al | Yes; Generator Integrated with Stand | 100 cm | contact manufacturer | <80 kg | AC, 180 - 260 V | Not applicable | unlimited exposures on AC | Not applicable | Not applicable | Not applicable | Wired | Temperature: 10 to 40 °C Humidity: 30 - 75% (non-condensing) Atmospheric pressure: 700 hPa - 1060 hPa | Temperature: -10 to 70°C Humidity: 5 - 95% (non-condensing) Atmospheric pressure: 500 - 1100 hPa | contact manufacturer |  | 43 x 35 cm (CareView® 1500Cwe) | 2 - 4 sec | 140 µm | 3.57 lp/mm | 71% at 0 lp/mm, 61% at 0.5 lp/mm 54% at 1 lp/mm, 46% at 1.5 lp/mm 40% at 2 lp/mm, 31% at 2.5 lp/mm 23% at 3 lp/mm, 16% at 3.5 lp/mm | contact manufacturer | AC (100 -250V) | Not applicable | continuous exposure | Not applicable | Not applicable | Not applicable | wireless | 3.7 kg | Yes DROC | contact manufacturer | contact manufacturer | IP64 for dust and water-resistance | contact manufacturer |  | Laptop/tablet: <3 Kg | LCD | English | Intel i5 | 8 GB | 1 TB | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer |  | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer |  | < 100 µSv/hr Measuring conditions: at 1 meter from Focal Spot | contact manufacturer | contact manufacturer |  | contact manufacturer | Yes | Yes |  | Yes | Yes | Yes | Yes 2 Years warranty | Contact manufacturer | Contact manufacturer |  | Contact manufacturer | Contact manufacturer |  | Contact manufacturer | Contact manufacturer | contact manufacturer
ALERIO | India | https://www.alerio.in/ | sales@alerio.in office@iatome.in | contact manufacturer | ALERIO Smart 8000 | Yes | 120 kg approx. (X-ray generator only) | Already assembled |  | 8 kW | 40 -120 kV (1 kV steps) | 1 - 250 mAs | 110 mA | 0.01 - 5 sec | Stationary, 2.0 mm focal spot | 10.4 kHU | Yes; The unit is programmed with overload protection, an error message is generated during overload condition without the X-ray being exposed | Manual , built-in laser guided LED light | > 2.0 mm Al | Yes; Generator Integrated with Stand | 100 cm | contact manufacturer | <120 kg | AC, 150 - 260 V | Not applicable | unlimited exposures on AC | Not applicable | Not applicable | Not applicable | Wired | Temperature: 10 to 40 °C Humidity: 30 - 75% (non-condensing) Atmospheric pressure: 700 hPa - 1060 hPa | Temperature: -10 to 70°C Humidity: 5 - 95% (non-condensing) Atmospheric pressure: 500 - 1100 hPa | contact manufacturer |  | 43 x 35 cm (CareView® 1500Cwe) | 2 - 4 sec | 140 µm | 3.57 lp/mm | 71% at 0 lp/mm, 61% at 0.5 lp/mm 54% at 1 lp/mm, 46% at 1.5 lp/mm 40% at 2 lp/mm, 31% at 2.5 lp/mm 23% at 3 lp/mm, 16% at 3.5 lp/mm | contact manufacturer | AC (100 -250V) | Not applicable | continuous exposure | Not applicable | Not applicable | Not applicable | wireless | 3.7 kg | Yes DROC | contact manufacturer | contact manufacturer | IP64 for dust and water-resistance | contact manufacturer |  | Laptop/tablet: <3 Kg | LCD | English | Intel i5 | 8 GB | 1 TB | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer |  |  | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer |  | < 100 µSv/hr Measuring conditions: at 1 meter from Focal Spot | contact manufacturer | contact manufacturer |  | contact manufacturer | Yes | Contact manufacturer |  | Yes | Yes | Yes | Yes 2 Years warranty | Contact manufacturer | Contact manufacturer |  | Contact manufacturer | Contact manufacturer |  | Contact manufacturer | Contact manufacturer | contact manufacturer
ASPEN IMAGING | United States | https://aspen-imaging.com/europa.html | sales@aspenstate.com |  | Europa | Yes | 7.25 kg | contact manufacturer |  | contact manufacturer | 60 kV (Fixed) | 0.06 -1.6 mAS | 2 mA (Fixed) | 0.03 - 1.3 sec | Stationary (contact manufacturer for focal spot size) | contact manufacturer | Yes; the error code prohibits the exposure | Manual shutter type, lead | 0.5 mm AI | Yes; separate component | 30 cm | No | contact manufacturer | Battery, 12.2 VDC | 1.5 hours | 100 exposures | Yes | No, available upon request | Yes | Wired ( 6 m hand switch, deadman type) | Temperature: -10 to 55°C Humidity: 5 - 80% (non-condensing) | Temperature: 20 to 40°C Humidity: 30 - 75% Atmospheric pressure: 150 - 500 hPa | No |  | 35 x 43 cm (DRTECH EVS3643W) | < 7 sec | 140 µm | contact manufacturer | ≥ 45% at 0.5 lp/mm | contact manufacturer | Battery | AC (110 V) | contact manufacturer | 1.5 hours | Yes | Yes (1 extra battery) | wired or wireless | 3.2 kg | Yes EConsole1 | Temperature : -15 to 55 °C Humidity : 10 - 90% Atmospheric pressure : 700 - 1060 hPa | Temperature: 10 to 35°C Humidity: 30 - 85% Atmospheric pressure : 700 - 1060 hPa | IPx4, 1 m drop test | Yes; optional |  | Laptop/tablet: 1.5 kg | 1920x1080 | English; Spanish (Windows dependent) | i7 12th gen or higher | contact manufacturer | 1 TB | wired or wireless transfer | Temperature: -15 to 55°C Humidity: 10 - 90% Atmospheric pressure: 500 to 1060 hPa | Temperature: 10 to 35°C Humidity: 30 - 85% Atmospheric pressure: 700 - 1060 hPa | contact manufacturer |  | Yes | Yes | contact manufacturer | contact manufacturer | Yes | No | Optional; pelican card case |  | contact manufacturer | Negotiable | Negotiable |  | FDA 510k | Yes | Yes |  | Yes | Yes | Yes | Yes | Yes | Yes |  | Yes | Yes |  | contact manufacturer | No | contact manufacturer
ASPEN IMAGING | United States | https://aspen-imaging.com/apx-24.html | sales@aspenstate.com |  | APX | Yes | 21.3 kg | contact manufacturer |  | 2.4 kW | 40 to 100 kV (1 kV steps) | 0.32 to 10 mAs | 40 mA | 0.01 - 2 sec | Stationary, 2.0 mm focal spot | contact manufacturer | Yes; the LED indicator if activated if the system is overheated or not ready for exposure | Manual shutter type, lead | 2.0 mm AI | Yes; separate component | 30 cm | No | contact manufacturer | AC, 110 - 220 V, 50/60 Hz | Not applicable | unlimited exposures on AC | Not applicable | Not applicable | Not applicable | Wired ( 6 m hand switch, deadman type) | Temperature: -25 to 70°C Humidity: 5 - 95% (non-condensing) | Temperature: 10 to 40°C, Humidity: 30% - 75%, Atmospheric pressure: 110 - 700 hPa | No |  | 35 x 43 cm (DRTECH EVS3643W) | < 7 sec | 140 µm | contact manufacturer | ≥ 45% at 0.5 lp/mm | contact manufacturer | Battery | AC (110 V) | contact manufacturer | 1.5 hours | Yes | Yes (1 extra battery) | wired or wireless | 3.2 kg | Yes EConsole1 | Temperature : -15 to 55 °C Humidity : 10 - 90% Atmospheric pressure : 700 - 1060 hPa | Temperature: 10 to 35°C Humidity: 30 - 85% Atmospheric pressure : 700 - 1060 hPa | IPx4, 1 m drop test | Yes; optional |  | Laptop/tablet: 1.5 kg | contact manufacturer | English; Spanish (Windows dependent) | i7 12th gen or higher | contact manufacturer | 1 TB | wired or wireless trasfer | Temperature: -15 to 55°C Humidity: 10 - 90% Atmospheric pressure: 500 to 1060 hPa | Temperature: 10 to 35°C Humidity: 30 - 85% Atmospheric pressure: 700 - 1060 hPa | contact manufacturer |  | Yes | Yes | contact manufacturer | Yes; contact manufacturer for details | Yes | No | Optional; pelican hard case |  | contact manufacturer | Negotiable | Negotiable |  | FDA 510k | contact manufacturer | contact manufacturer |  | Yes | Yes | Yes | Yes | Yes | Yes |  | Yes | Yes |  | contact manufacturer | No | contact manufacturer
DELFT IMAGING | Netherlands | https://dc.delft.care/x-ray-systems/ | info@delft.care |  | Delft Light | Yes | 25 kg | 10 minutes |  | 1.35 kW | 40-90 kV (2kV steps) | 0.01 - 20 mAs | 15 mA | 0.01-1.0 sec | Stationary, Tungsten Anode, 0.8 mm focal spot | 10 kHU | Yes; Error shown in case of overheating - the system will not expose | Manual multi leaf with LED light, Dual laser positioning pointer: Class III-a laser diode | 2.5 mm Al | Yes; separate component. Multiple options, among which a foldable, rotatable, adjustable stand with balanced spring-arm | Adjustable SID, recommended minimally 120 cm | Yes; Option available | 7 kg | Battery, 59.2 VDC | 2 - 3 hours | > 200 exposures | No | Yes, upon request | Yes | Wired exposure switch | Temperature: 10 to 60°C Humidity: 10 - 90% (non-condensing) Atmospheric pressure: 500-1060 hPa | Temperature: 10 to 40°C Humidity: 30 - 75% (non-condensing) Atmospheric pressure: 700-1060 hPa | contact manufacturer |  | 36 x 43 cm | < 3 sec | 140 μm | > 4.0 lp/mm | 85% | contact manufacturer | Battery, Li-ion | AC, Battery charger | 200 + | 2 hours | No | Yes; extra battery included | wired and wireless | 2.9 kg | Yes Canon CXDI NE | contact manufacturer | Temperature: 5 to 35°C Humidity: 30 - 80% (non-condensing) Atmospheric pressure: 700 - 1060 hPa | IP55 certification against limited dust ingress and water protection | Yes VersariX adjustable detector hanger |  | Laptop | Full HD, 16:9, 1920x1080 pixels | Multi language | Intel | 8 GB | 500 GB SSD | via DICOM or local network connections | contact manufacturer | contact manufacturer | contact manufacturer |  | Yes | Yes | Yes; Manual | Yes; auto processing | Yes | Yes; CAD4TB | Yes; Light Backpack |  | contact manufacturer | contact manufacturer | contact manufacturer |  | CE certification | Yes | contact manufacturer |  | Yes | Yes | Yes | Yes 1 year after training/15 months after shipment (whichever comes first) | contact manufacturer | contact manufacturer |  | Yes; extensive experience in organizing training and installation sessions (online or onsite), support office located in Accra, Ghana where all the topics about safety, utilization, transportation, maintenance, calibration, and more are covered | Yes; can be discussed based on project needs |  | Yes Extended warranty options, The warranty extension packages include updates, upgrades, software pactches, spare part provision (Excluding batteries) transported via DAP incoterm and 24/7 helpdesk support | Yes Significant experience with TB programs in LMICs, presence in over 70 countries - with local service coverage in each country, the solutions have screened over 18 million people worldwide. Delft Imaging is a total solution provider, offering X-ray, AI, training and after-sales services under one roof. | contact manufacturer
DELFT IMAGING | Netherlands | https://dc.delft.care/about/ | info@delft.care |  | Delft Ultra | Yes | 15 kg | 5 minutes |  | 0.45 kW | 50-90 kV (1kV steps) | 0.04 - 6.5 mAs | 5 mA | 0.02-1.30 sec | Stationary, Tungsten Anode, 0.8 mm focal spot | 5 kHU | Yes; Error shown in case of overheating - the system will not expose | Manual mutli leaf with LED light | 2.5 mm AI | Yes; separate component. Multiple options, among which a tripod with spindle | Adjustable SID, recommended minimally 100 cm | Yes; Option available | 5.2 kg | Battery, 25.2 VCD 1.5 A AC (100-240 V, 50/60 Hz, 1.5A) for charging | 1.5 - 2 hours | > 100 exposures | Yes | No | Yes | Wired exposure switch | Temperature: - 20 to 50°C Humidity: 10 - 90% (non-condensing) Atmospheric pressure: 500-1060 hPa | Temperature: 10 to 30°C, Humidity 30 - 75% (non-condensing) Atmospheric pressure: 700-1060 hPa | contact manufacturer |  | 36 x 43 cm | 2 sec | 140 μm | > 3.5 lp/mm | 75% | contact manufacturer | Battery, Li-ion | AC, Battery charger | 200 + | 2 hours | Yes | Yes; extra battery included | wired and wireless | 3 kg | Yes VXvue / Delft Imaging | contact manufacturer | contact manufacturer | contact manufacturer | Yes VersariX adjustable detector hanger |  | Laptop | Full HD, 16:9, 1920x1080 pixels | Multi language | Intel | 8 GB | 500 GB SSD | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer |  | Yes | Yes | Yes; Manual | Yes; auto processing | Yes | Yes; CAD4TB - In the future this will likely be fully integrated | Yes; contact manufacturer for details |  | contact manufacturer | contact manufacturer | contact manufacturer |  | FDA pending | contact manufacturer | contact manufacturer |  | Yes | Yes | contact manufacturer | Yes 1 year after training, 15 months after shipment (whichever comes first) | contact manufacturer | contact manufacturer |  | Yes; extensive experience in organizing training and installation sessions (online or onsite), support office located in Accra, Ghana where all the topics about safety, utilization, transportation, maintenance, calibration, and more are covered | Yes; can be discussed based on project needs |  | Yes Extended warranty options, The warranty extension packages include updates, upgrades, software pactches, spare part provision (Excluding batteries) transported via DAP incoterm and 24/7 helpdesk support | Yes Significant experience with TB programs in LMICs, presence in over 70 countries - with local service coverage in each country, the solutions have screened over 18 million people worldwide. Delft Imaging is a total solution provider, offering X-ray, AI, training and after-sales services under one roof. | contact manufacturer
ECORAY | South Korea | http://ecoray.kr/pg_acpowered#bp | info@ecoray.kr |  | mini90 | Yes | 7.5 kg | 1 to 3 minutes |  | 0.8 kW | 40 - 90 kV (1kV steps) | 0.32 - 20 mAs | 10 mA | 0.04 - 2 sec | Stationary, 0.8 mm focal spot | 11.2 kHU | Yes; Program protects overload combination like kV, mAs and unit produces heat error code and prevents exposure until the generator cools down | Manual, lead shutter, LED lighting | 3.2 mm Al | Yes; separate component | flexible, up to 180 cm | contact manufacturer | 4 kg | Battery | 1.3 hours | 150 exposures | Yes | No | Yes | Wired and remote exposure switch | Temperature: -20 to 45°C Humidity: 10 - 95% Atmospheric pressure: 500 - 1060 hPa | Temperature: 10 to 40 °C Humidity: 30 - 75 % Atmospheric pressure: 700 - 1060 hPa | contact manufacturer |  | 35 x 43 cm | < 4 sec | 140 μm | > 3.5 lp/mm | 46% at 0.5lp/mm | contact manufacturer | Battery | contact manufacturer | contact manufacturer | 2.5 hours | Yes | No | wireless | 3.15 kg | contact manufacturer | Temperature: -15 to 55°C Humidity: 5 - 90% (non-condensing) Atmospheric pressure: 500 - 1060hPa | Temperature: 0 to 40°C Humidity: 5 - 90% (non-condensing) Atmospheric pressure: 700-1060 hPa | Factory tested for 200 Kg physical shock test Drop test at 1m height IP67 for water/dust penetration Contact manufacturer for details | No |  | Laptop (14 to 16") | contact manufacturer | English | i7 (Pentium) | 8 GB | 512 GB | via DICOM network or manual file transfer by USB | Temperature: 0 to 40°C Humidity: 5 - 90% (non-condensing) Atmospheric pressure: 700 - 1060hPa | Temperature: 2 to 40°C Humidity: 5 - 90% (non-condensing) Atmospheric pressure: 700 - 1060hPa | Yes: matches MIL-STD-810H test |  | Yes | Yes | contact manufacturer | Yes (automatic/manual) | Yes | Yes; contact manufacturer for details | contact manufacturer |  | contact manufacturer | No | No |  | FDA 510k is in the process (will be obtained in 2024) | Yes | contact manufacturer |  | Yes | Yes | Yes | Yes 1 year from the date of shipment | Yes | Yes; in User manual |  | Optional; depends on contract | Optional; depends on contract |  | Yes contact manufacturer for details | No | contact manufacturer
ECORAY | South Korea | http://ecoray.kr/pg_acpowered#bp | info@ecoray.kr |  | 90BT | Yes | 12.2 kg | 1 to 3 minutes |  | 1.35 kW | 40 - 90 kV (1 kV steps) | 0.32 - 30 mAs | 18 mA | 0.016 - 2 sec | Stationary, 1.2 mm focal spot | 14 kHU | Yes; Program protects overload combination like kV, mAs and unit produces heat error code and prevents exposure until the generator cools down | Manual, lead shutter, LED lighting | 2.59 mm Al | Yes; separate component. Wheel stand available with brake (Aluminum made) | flexible, up to 180 cm | contact manufacturer | 8.7 kg | Battery | 3.5 hours | 350 exposures | Yes | No | Yes | Wired (3 m hand switch) | Temperature: -20 to 45°C Humidity: 10 - 95% Atmospheric pressure: 500 - 1060 hPa | Temperature: 10 to 40 °C Humidity: 30 - 75 % Atmospheric pressure: 700 - 1060 hPa | contact manufacturer |  | 35 x 43 cm | < 4 sec | 140 μm | > 3.5 lp/mm | 46% at 0.5lp/mm | contact manufacturer | Battery | contact manufacturer | contact manufacturer | 2.5 hours | Yes | No | wireless | 3.15 kg | contact manufacturer | Temperature: -15 to 55°C Humidity: 5 - 90% (non-condensing) Atmospheric pressure: 500 - 1060hPa | Temperature: 0 to 40°C, Humidity: 5 - 90% (non-condensing) Atmospheric pressure: 700-1060 hPa | Factory tested for 200 Kg physical shock test Drop test at 1m height IP67 for water/dust penetration Contact manufacturer for details | No |  | Laptop (14 to 16") | contact manufacturer | English | i7 (Pentium) | 8 GB | 512 GB | via DICOM network or manual file transfer by USB | Temperature: 1 to 40°C Humidity: 5 - 90% (non-condensing) Atmospheric pressure: 700 - 1060hPa | Temperature: 2 to 40°C Humidity: 5 - 90% (non-condensing) Atmospheric pressure: 700 - 1060hPa | Yes: matches MIL-STD-810H test |  | Yes | Yes | contact manufacturer | Yes (automatic/manual) | Yes | Yes; contact manufacturer for details | contact manufacturer |  | contact manufacturer | No | No |  | contact manufacturer | Yes | contact manufacturer |  | Yes | Yes | Yes | Yes | Yes | Yes; in User manual |  | Optional; depends on contract | Optional; depends on contract |  | contact manufacturer | No | contact manufacturer
ECORAY | South Korea | http://ecoray.kr/pg_acpowered#bp | info@ecoray.kr |  | 9020BT | Yes | 13.3 kg | 1 to 3 minutes |  | 1.8 kW | 41 - 90 kV (1kV steps) | 0.32 - 40 mAs | 25 mA | 0.01 - 2.6 sec | Stationary, 0.8 mm focal spot | 9.8 kHU | Yes; Program protects overload combination like kV, mAs and unit produces heat error code and prevents exposure until the generator cools down | Manual, lead shutter, LED lighting | 2.89 mm Al | Yes; separate component. Wheel stand available with brake (Aluminum made) | flexible, up to 180 cm | contact manufacturer | 9.8 kg | Battery | 3.5 hours | 350 exposures | Yes | No | Yes | Wired (3 m hand switch) | Temperature: -20 to 45°C Humidity: 10 - 95% Atmospheric pressure: 500 - 1060 hPa | Temperature: 10 to 40 °C Humidity: 30 - 75 % Atmospheric pressure: 700 - 1060 hPa | contact manufacturer |  | 35 x 43 cm | < 4 sec | 140 μm | > 3.5 lp/mm | 46% at 0.5lp/mm | contact manufacturer | Battery | contact manufacturer | contact manufacturer | 2.5 hours | Yes | No | wireless | 3.15 kg | contact manufacturer | Temperature: -15 to 55°C Humidity: 5 - 90% (non-condensing) Atmospheric pressure: 500 - 1060hPa | Temperature: 0 to 40°C Humidity: 5 - 90% (non-condensing) Atmospheric pressure: 700-1060 hPa | Factory tested for 200 Kg physical shock test Drop test at 1m height IP67 for water/dust penetration Contact manufacturer for details | No |  | Laptop (14 to 16") | contact manufacturer | English | i7 (Pentium) | 8 GB | 512 GB | via DICOM network or manual file transfer by USB | Temperature: 2 to 40°C Humidity: 5 - 90% (non-condensing) Atmospheric pressure: 700 - 1060hPa | Temperature: 2 to 40°C Humidity: 5 - 90% (non-condensing) Atmospheric pressure: 700 - 1060hPa | Yes: matches MIL-STD-810H test |  | Yes | Yes | contact manufacturer | Yes (automatic/manual) | Yes | Yes; contact manufacturer for details | contact manufacturer |  | contact manufacturer | No | No |  | contact manufacturer | Yes | contact manufacturer |  | Yes | Yes | Yes | Yes | Yes | Yes; in User manual |  | Optional; depends on contract | Optional; depends on contract |  | contact manufacturer | No | contact manufacturer
EPSILON | India | https://www.epsilonhealthcare.com/portablehf.php | marketing@epsilonhealthcare.com |  | EP Corsa 2.4 P | Yes | < 20 kg (X-ray generator only) | contact manufacturer |  | 2.4 kW | 40 – 100 kV (1 kV steps) | contact manufacturer | 35 mA | contact manufacturer | Stationary, single focus (contact manufacturer for size) | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | Battery | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer |  | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer |  | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer |  | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer |  | contact manufacturer | contact manufacturer | contact manufacturer |  | contact manufacturer | contact manufacturer | contact manufacturer |  | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer |  | contact manufacturer | contact manufacturer |  | contact manufacturer | contact manufacturer | contact manufacturer
FUJIFILM | Japan | https://www.fujifilm.com/rs/en/healthcare/x-ray/portable-x-ray/fdr-xair | Contact FIND for contact info |  | FDR Xair | Yes | 25 kg | 10 minutes |  | 0.45 kW | 50-90 kV (2 kV steps) | 0.2 - 2.5 mAs | 5 mA | 0.04-0.5 sec | Stationary, 0.8 mm focal spot | 10 kHU | Yes; An error sign will appear | Built-in LED Collimator | 2.5 mm Al | Yes; separate component. lightweight (2 kg) verticle stand | 100cm | Yes | 3.5 kg | Battery | 4.5 hours | 100 exposures (at 90kV, 0.5mAs with power bank) | Yes | Yes | Yes, 5 battery bar indicator available | Wired and remote exposure switch | Temperature: -10 to 50°C Humidity: 10 - 90% Atmospheric pressure: 700-1060 hPa | Temperature: 15 to 30°C Humidity: 15 - 80% Atmospheric pressure: 700-1060 hPa | Yes Conducted Shock/Drop test. According to the degree of protection against harmful ingress of water IPX0. If any dust adheres to the operation panel, wipe it off with a soft cloth moistened with water. |  | FDR D-EVO II 35 x 43 cm, or FDR D-EVO II 43 x 43 cm | 3 sec | 100 µm | 1 lp/mm | 45% for GoS scintillator detector 72% for CsI scintillator detector | contact manufacturer | Battery | AC (100 to 240 V, 50-60Hz) | contact manufacturer for exposure numbers 4 hours (equipped 2 batteries, total 8 hours of operation) | approx. 3 hours | No | Yes | wired and wireless | 2.6 kg | Yes Console Advance Image Processing from FujiFilm, includes virtual grid and dynamic visualization software | Temperature: -10 to 50°C Humidity: 10 - 90% Atmospheric pressure: 700-1060 hPa | Temperature: 10 to 30°C Humidity: 15 - 90% Atmospheric pressure: 700-1060 hPa | Factory tested for 350 Kg physical shock test IPX6 waterproofing contact manufacturer for details | Yes Lightweight at 2kg. Vertical movements. Two-way height adjustment can be done based on the two-section splittable center. First height adjustment can be done with leg extention, and the second can be done by elevation function of the center column, performed by coiling the knob. No wheels but it is lightweight and easy to move. |  | Laptop, 1.5kg | contact manufacturer | 21 languages (English, French, Spanish, Germany, Chinese, Russia, others) | Intel® Core™ i5-8265U, 4 core, 1.6-3.9GHz processor | 8 GB | 256 GB S-ATA | via DICOM network | Temperature: -10 to 50°C Humidity: 5 - 95% Atmospheric pressure: 700-1060 hPa | Temperature: 5 to 35°C Humidity: 15 - 80% (non-condensing) | Yes; matches MIL-STD810 test |  | Yes | Yes | Yes; Manual | Yes clip, zoom, magnifier, invert, rotate, flip, annotations, measurements, digital collimation, etc.), image view, detail enhancement and noise suppression, tissue equalization, alphanumeric annotation of images | Yes | Yes; Lunit Insight, Qure.AI qXR, Delft and Infervision | Yes; suitcase |  | 0.0134µSv (at 1m), 0.0053µSv (at 2m) Measuring condition: X-ray forward radiation, 90kV, 2.5mAs, collimator closed | Yes; 0.5mm Pb Equivalent | No |  | CE-IVDD, PMDA | Yes | Yes |  | Yes | Yes | Yes | Yes 1 year from installation by manufacturer warranty | Yes | Yes |  | Yes; 1 day training | Yes; flexible training program and locations |  | Yes Extended warranty (option) Can choose upto 4 years (2nd to 5th year). The warranty contract covers: -1 time preventive maintenance per year -Machine (including X-RAY TUBE) -Mobile Console -FPD DEVOII FPD | Yes | contact manufacturer
FUJIFILM | Japan | https://www.fujifilm.com/ca/en/healthcare/digital-radiography/digital-x-ray/mobile-dr-systems/fdr-nano | Contact FIND for contact info |  | FDR nano | Yes | 90 kg | 10 minutes |  | 2.5 kW | 40-100 kV | 0.25 - 25 mAs | 25 mA | 0.01-1 sec | Stationary, 1.2 mm focal spot | 50 kHU | Yes; An error sign will appear | Multileaf collimator with positioning light | 1.2 mm Al | Yes; built into the system | 100cm | Yes | 90 kg (integrated weight of equipment itself) | Battery | 4 hours 1 hour use time with 15 mins of charge | 240 exposures (20 exposures/hr and up to 12 hours of use with full charge) | Yes | No | Yes, 5 battery bar indicator available | contact manufacturer | Temperature: -10 to 50°C Humidity: 10 - 90% Atmospheric pressure: 700-1060 hPa | Temperature: 15 to 30°C Humidity: 15 - 80% Atmospheric pressure: 700-1060 hPa | contact manufacturer |  | FDR D-EVO II 35 x 43 cm, or FDR D-EVO II 43 x 43 cm | < 3 sec | 100 µm | 1 lp/mm | 45% for GoS scintillator detector 72% for CsI scintillator detector | contact manufacturer | Battery | AC (100 to 240 V, 50-60Hz) | contact manufacturer for exposure numbers 4 hours (equipped 2 batteries, total 8 hours of operation) | approx. 3 hours | No | Yes | wired and wireless | 2.6 kg | Yes Console Advance Image Processing from FujiFilm, includes virtual grid and dynamic visualization software | Temperature: -10 to 50°C Humidity: 10 - 90% Atmospheric pressure: 700-1060 hPa | Temperature: 10 to 30°C, Humidity: 15 - 90% Atmospheric pressure: 700-1060 hPa | Factory tested for 350 Kg physical shock test IPX6 waterproofing contact manufacturer for details | Yes Lightweight at 2kg. Vertical movements. Two-way height adjustment can be done based on the two-section splittable center. First height adjustment can be done with leg extention, and the second can be done by elevation function of the center column, performed by coiling the knob. No wheels but it is lightweight and easy to move. |  | Monitor integrated with the system | Integrated console panel with multitouch screen | 21 languages (English, French, Spanish, Germany, Chinese, Russian, others) | 4th Gen Intel® Core™ i5-4300U, 2-Core, 1.9GHz processor | 8 GB | 60 GB SSD | via DICOM network | Temperature: -10 to 50°C Humidity: 5 - 95% Atmospheric pressure: 700-1060 hPa | Temperature: 15 to 30°C Humidity: 15 - 80% Atmospheric pressure: 700-1060 hPa | Contact manufacturer |  | Yes | Yes | Yes; contact manufacturer for details | Yes clip, zoom, magnifier, invert, rotate, flip, annotations, measurements, digital collimation, etc.), image view, detail enhancement and noise suppression, tissue equalization, alphanumeric annotation of images | Yes | Yes; Lunit Insight | No |  | 6.7µSv (at 1m), 1.8µSv (at 2m) Measuring condition: X-ray forward radiation, 80kV, 12.5mAs, measurement height: 350mm | Yes | No |  | CE-IVDD, PMDA, Brazil and FDA 510k | Yes | Yes |  | Yes | Yes | Yes | Yes 1 year from installation by manufacturer warranty | Yes | Yes |  | Optional; depends on contract | Yes; flexible training program and locations |  | Yes contact manufacturer for details | No | contact manufacturer
KE MEDITEC | Germany | http://ke-meditec.com/index.php/en/atx-quantpower-400-dr-eng-menue | info@ke-meditec.com |  | ATX QUANTpower 400 DR | Yes | 75 kg (X-ray generator only) | contact manufacturer |  | 6 kW | 40 - 120 kV (2kV steps) | 0.4 - 143 mAs | 100 mA | 0.01 – 2 sec | Stationary, 1.2 mm focal spot | contact manufacturer | Yes; auto programmed | Manual, rotating shuttere LED light | 3.2 mm Al | Yes; Generator Integrated with Stand, wheeled base | contact manufacturer | contact manufacturer | contact manufacturer | AC, 200 - 260 V, 50 - 60Hz | Not applicable | unlimited exposures on AC | Not applicable | Not applicable | Not applicable | contact manufacturer | contact manufacturer | contact manufacturer | Yes; Splashproof (according to IEC specification 529) |  | 43 x 35 cm (Canon) | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer |  | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer |  | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer |  | contact manufacturer | contact manufacturer | contact manufacturer |  | contact manufacturer | contact manufacturer | contact manufacturer |  | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer |  | contact manufacturer | contact manufacturer |  | contact manufacturer | contact manufacturer | contact manufacturer
MED SING LONG | China | https://www.medicalequipment-msl.com/htm/portable-x-ray-machine/Portable-High-Frequency-X-ray-Machine-MSLGX08.html | cindy@medicalequipment-msl.com |  | MSLGX08 | Yes | < 20 kg (X-ray generator only) | contact manufacturer |  | 2.8 / 3.5 / 5.6 / 8.0 kW options | 40 – 125 kV | 0.1 - 12.5 mAs / 0.1 - 320 mAs/ 0.1 - 360 mAs options | contact manufacturer | 0.01 - 10 sec | Stationary, single focal spots 1.0 / 1.5 / 1.8 / 2.0 mm options | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | Battery | 2 hours / 4 hours for higher power versions | 100 exposures / 500 exposures for higher power versions | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer |  | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer |  | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer |  | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer |  | contact manufacturer | contact manufacturer | contact manufacturer |  | contact manufacturer | contact manufacturer | contact manufacturer |  | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer |  | contact manufacturer | contact manufacturer |  | contact manufacturer | contact manufacturer | contact manufacturer
Medical ECONET | Germany | https://www.medical-econet.com/en/digital-radiography/generators/106/mex-100-human-version?c=61 | info@medical-econet.com |  | meX+100 | Yes | 19.6 kg (X-ray generator only) | contact manufacturer |  | 5 kW (at 75 kV) | 40 - 110 kV (1 kV steps) | 0.1 - 100 mAs | 100 mA | contact manufacturer | Stationary, 1.5 mm focal spot | 76 kHU | contact manufacturer | Manual LED, dual laser pointer max light field 35 x 35 cm at 65 cm rotating collimator optional | contact manufacturer | optional as a separate component | contact manufacturer | contact manufacturer | contact manufacturer | AC, 100 - 240 V(Free-Voltage), 50 - 60 Hz | Not applicable | unlimited exposures on AC | Not applicable | Not applicable | Not applicable | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer |  | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer |  | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer |  | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer |  | contact manufacturer | contact manufacturer | contact manufacturer |  | contact manufacturer | contact manufacturer | contact manufacturer |  | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer |  | contact manufacturer | contact manufacturer |  | contact manufacturer | contact manufacturer | contact manufacturer
MINXRAY | United States | https://www.minxray.com/medical | info@minxray.com |  | Impact | Yes | 20 kg (approx.) | 5 minutes |  | 1.35 kW (at 90 kV) | 40 to 90 kV (2 kV steps) | contact manufacturer | 15 mA | 0.02 – 1.0 sec | Stationary, 0.8 mm focal spot | 9.86 kHU | Yes; System has fixed duty cycle that prevents exposures. It ensures maximum life of the x-ray tube and avoids interruption of operation, System will not exceed the duty cycle of 1:60. (For example, after a 1 second exposure, allow 60 seconds before initiating the next exposure.) | Continuously adjustable light beam type with adjustable dual integrated laser pointers to indicate central ray and distance in bright ambient light conditions LED light; Class II laser diode, max output ≤1.0 mW, wavelength 635 nm | 2.0 mm Al | Yes; X300S Tripod Maximum Working Height 66.9 inches (169.9 cm) Minimum Working Height 3.5 inches (8.9 cm) Closed Length 24 inches (61 cm) Weight 5 lbs (2.5 kg) Box Dimensions 25.5 (L) x 5.5 (W) x 5.3 (H) inches (64.8 x 14.0 x 15.5 cm) | 180 cm (72 inches) | Yes:; angle meter available | 7.4 kg | Battery AC (Single phase 90 - 240 V, 50/60Hz) for charging | Approx. 3 - 4 hours | 200 exposures (at 90 kVp at 1 mAs) | No | Yes | Yes | Wired (7 m hand switch) | Temperature: -20 to 50°C Humidity: max 85% (non-condensing) | Temperature: -20 to 50°C Humidity: max 85% (non-condensing) | contact manufacturer |  | 43.0 x 35.6 cm | 6 sec | 140μm | 3.57 Ip/mm | 63% at 1.0 Ip/mm, 35% at 2.0 Ip/mm 17% at 3.0 Ip/mm, 10% at 3.5 Ip/mm 62% at 0.0 Ip/mm, 50% at 0.5 Ip/mm 45% at 1.0 Ip/mm), 38% at 1.5 Ip/mm 32% at 2.0 Ip/mm, 24% at 2.5 Ip/mm 17% at 3.0 Ip/mm, 10% at 3.5 Ip/mm | contact manufacturer | Battery | contact manufacturer | contact manufacturer Battery runtime of 3.5 to 4.5 hours per battery | 3 to 4 hours | Yes | Yes | wired or wireless | 3.7 kg | Yes D-XR | Temperature : 20 to 45°C Humidity : 15 - 95% | Temperature: 5 to 35°C Humidity : 30 - 75% | Yes, contact manufacturer for details | Yes Aluminum Tripod Maximum Working Height 66.9 inches (169.9 cm) Minimum Working Height 3.5 inches (8.9 cm) Closed Length 24 inches (61 cm) Weight 5 lbs (2.5 kg) Box Dimensions 25.5 (L) x 5.5 (W) x 5.3 (H) inches (64.8 x 14.0 x 15.5 cm) |  | Laptop | contact manufacturer | Multi language | Intel i7 or better | 64 GB | 1 TB SSD | wired or wireless transfer using DICOM | Temperature: -40 to 65°C Humidity: 0 - 95% (non-condensing) | Temperature: 0 to 35°C Humidity: 10% - 80% (non-condensing) | Yes,contact manyfacturer for details |  | Yes | Yes | Yes; contact manufacturer for details | Yes; auto/manual processing | Yes | Yes; Qure AI (online or offline implementation) | Yes; contact manufacturer for details |  | contact manufacturer | Yes; 0.35mm Pb Equivalent | No |  | FDA 510k | Yes | Yes |  | Yes | Yes | Yes | contact manufacturer | Yes | Yes |  | Yes | Yes |  | Yes contact manufacturer for details | Yes, systems in over 80 countries | contact manufacturer
MINXRAY | United States | https://www.minxray.com/medical | info@minxray.com |  | CMDR.CW.120.S | Yes | 65 kg (without carrying case), 115kg (including carrying case) | 3 minutes |  | 1.56 kW (at 120 kV) | 40 to 90 kV (2 kV steps) | contact manufacturer | 45 mA | 0.02 – 5.0 sec | Stationary, 1.2 mm focal spot | 65 kHU | Yes; System has fixed duty cycle that prevents exposures. It ensures maximum life of the x-ray tube and avoids interruption of operation, System will not exceed the duty cycle of 1:60. (For example, after a 1 second exposure, allow 60 seconds before initiating the next exposure.) | Continuously adjustable light beam type with adjustable dual integrated laser pointers to indicate central ray and distance in bright ambient light conditions LED light; Class II laser diode, max output ≤1.0 mW, wavelength 635 nm | 3.2 mm Al | Yes; XGSMK7 DR stand. Stainless steal. 16" solid wheels with 4 brakes. Stand uses gas shocks to list assist the tube. Max Height 206 CM | 180 cm (72 inches) | Yes; angle meter available | 17.5 kg | AC, 100-260 V; 50/60 Hz | Not applicable | unlimited exposures on AC | Not applicable | Not applicable | Not applicable | Wired ( 6.1 m hand switch) | Temperature: -20 to 50°C Humidity: max 85% (non-condensing) | Temperature: -20 to 50°C Humidity: max 85% (non-condensing) | contact manufacturer |  | 43.0 x 35.6 cm | 6 sec | 140μm | 3.57 Ip/mm | 63% at 1.0 Ip/mm, 35% at 2.0 Ip/mm 17% at 3.0 Ip/mm, 10% at 3.5 Ip/mm 62% at 0.0 Ip/mm, 50% at 0.5 Ip/mm 45% at 1.0 Ip/mm), 38% at 1.5 Ip/mm 32% at 2.0 Ip/mm, 24% at 2.5 Ip/mm 17% at 3.0 Ip/mm, 10% at 3.5 Ip/mm | contact manufacturer | AC (100- 240V, 50/60 Hz) | Not applicable | continuous exposure | Not applicable | Not applicable | Not applicable | wired or wireless | 3.7 kg | Yes D-XR | Temperature : 20 to 45°C Humidity : 15 - 95% | Temperature: 5 to 35°C Humidity: 30 - 75% | Yes, contact manufacturer for details | Yes Aluminum Tripod Maximum Working Height 66.9 inches (169.9 cm) Minimum Working Height 3.5 inches (8.9 cm) Closed Length 24 inches (61 cm) Weight 5 lbs (2.5 kg) Box Dimensions 25.5 (L) x 5.5 (W) x 5.3 (H) inches (64.8 x 14.0 x 15.5 cm) |  | Laptop | contact manufacturer | Multi language | Intel i7 or better | 64 GB | 1 TB SSD | wired or wireless transfer using DICOM | Temperature: -40 to 65°C Humidity: 0 - 95% (non-condensing) | Temperature: 0 to 35°C Humidity: 10% - 80% (non-condensing) | Yes, contact manufacturer for details |  | Yes | Yes | Yes; contact manufacturer for details | Yes; auto/manual processing | Yes | Yes; Qure AI (online or offline implementation) | Yes; contact manufacturer for details |  | contact manufacturer | Yes; 0.35mm Pb Equivalent | No |  | FDA 510k | Yes | Yes |  | Yes | Yes | Yes | contact manufacturer | Yes | Yes |  | Yes | Yes |  | Yes contact manufacturer for details | Yes, systems in over 80 countries | contact manufacturer
MYLAB | India | https://mylabglobal.com/ | business@mylabglobal.com |  | MyBeam | Yes | 15 kg (approx.) | 15 minutes |  | 0.21 kW | 40 - 70 kV (1 kV steps) | 0.3 - 6 mAs | 3 mA | 0.1 - 2 sec | Stationary, 0.4 mm focal spot | 0.60 kHU | Yes; the unit is programmed with overload protection, an error would get displayed on the unit during the situation without being exposed. | Fixed field, Laser light guided field indicator | ≥ 2.0 mm AI | Yes; separate component. Portable and lightweight tripod stand (1.4 kg) | 100 cm | contact manufacturer | <3.2 kg | Battery, 5 to 12 VDC, 2.1 A AC (90-264 V, 50/60Hz, 3A) for charging | 0.5 hours | 250 exposures | No | No | Yes | Wired and remote controlled hand switch | Temperature: 0 to 50°C Humidity: max 70% Atmospheric pressure: 500 - 1060 hPa | Temperature: 0 to 40°C Humiditiy: 70% (max) Atmospheric pressure: 500 - 1060 hPa | contact manufacturer |  | 35 x 43 cm | 2 sec | 140 µm | 3.57 lp/mm | 72% at 0.1 lp/mm | contact manufacturer | AC (100 -240V, 47-63Hz) DC (24 V, 2.1 A) | contact manufacturer | contact manufacturer | contact manufacturer | Not applicable | Not applicable | wired or wireless | 3.4 kg | Yes LG DXD | Temperature : -20 to 60°C Humidity : 0 - 90% (non-condensing) Atmospheric Pressure : 500 - 1060 hPa | Temperature: -20 to 60°C Humidity: 0 - 90% (non-condensing) Atmospheric pressure : 500 - 1060 kPa | IP 53 for water and dust penetration The detector will be protected from dust and water splash at angles < 60 degree. | Yes, detector stand is provided for taking CXR in erect position |  | Laptop, 1.8 kg | 39.6 cm (15.6") diagonal, FHD (1920 x 1080), IPS, narow bezel, anti-glare, 250 nits, 45% NTSC | English | 12th Generation Intel Core i5 processor | 16 GB | 512 GB | wireless or manual transfer via USB ports | contact manufacturer | contact manufacturer | contact manufacturer |  | Yes | Yes | Yes; Manual | Yes; auto/manual processing | Yes | Yes AI integration for CXR TB screening AI Processing type : Offline AI Processing Time : max 2 min, Chest Pathologies covered : Tuberculosis, Normal-Abnormal (Pathologies) Contact manufacturer for additional details | Yes; cushioned/padded case or backpack |  | < 100 µSv/hr Measuring conditions: at 1 meter from Focal Spot | Yes | Yes; Thyroid collar apron is included |  | Certified by AERB (Atomic Energy Regulatory Board), India IS7620 BIS approval | Yes | No |  | Yes | Yes | Yes | Yes 1 year of warranty from date of billing | Yes | Yes |  | Yes | Yes |  | Yes Standard warranty is 1 year from the date of billing | Yes | contact manufacturer
OR TECHNOLOGY | Germany | https://www.oehm-rehbein.de/de | info@or-technology.com |  | Amadeo M-DR mini | Yes | 90 kg (approx.) | less than 1 minute |  | 5.0 kW | 40 - 110 kV (1 kV steps) | 0.1 - 100 mAs | 50 mA | contact manufacturer | Stationary, Tungsten Anode, 1.5 mm focal spot | 76 kHU | Yes; through Capacitor | Manual shutter with high brightness and integrated timer dual laser pointer Min light field : 5 cm x 5 cm at 100 cm Max. light field : 47 cm x 47 cm at 100 cm Lamp : 3.3 V, LED Auto timer : 30 sec | ≥ 3.0 mm Al | Yes; built into the system | 100 cm | Yes | Approx. 20.5 kg | AC, 100 - 240 VAC Capacitor back up at limited capacity | no battery available , capacitor backup for limited use, capacitor charging while in use | Unlimited exposures wit AC up to 20 exposure while used on fully charged capacitor | Not applicable No battery available, capacitor charging while in use | Not applicable | Yes, for the capacitor | Wired, two-stage hand trigger switch | Temperature: -10 to 55°C Humidity: 10 - 90% Atmospheric pressure: 700 - 1060 hPa If the unit was transported or stored at temperatures under or above the recommended temperature range, a warm up or cool down period must be observed prior to operation of the unit | Temperature: 10 to 30 °C Humidity: 50 - 70 % Atmospheric pressure: 700 - 1060 hPa | Yes; Protection against electric shock class 1 type B under IEC 60601-1 standard |  | 35 x 43 cm | 3 sec | 100 µm | 4.3 lp/mm | 73.4% at 0 lp/mm | contact manufacturer | Battery (24 V) | AC (Battery charger including 2x batteries) | contact manufacturer for exposure numbers battery runtime of 8 hours | approx. 4 hours | No | Yes | wireless | 3.0 kg | Yes dicomPACS DX-R | Temperature: -20 to 55 °C Humidity: 5 - 95 % (non-condensing) Atmospheric pressure: 600 - 1060 hPa Altitude: max. 3000 meters | Temperature: -20 to 55°C Humidity: 5 - 95 % (non-condensing) Atmospheric pressure: 600 - 1060 hPa Altitude: max. 3000 meters | contact manufacturer | Yes • Mobile special stand for DR detector / CR cassette and grid (35x43cm) • Available for different cassette formats, please specify detector type (manufacturer model) when ordering • Very flexible positioning of cassette and grid (due to vertical, horizontal and swivel joints as well as integrated spirit level) • Weight: 20 kg |  | Monitor integrated within the system, 2 kg | 1920 x 1080 resolution, 15.6" (39.6 cm) Multi-Touch Display | Multi language | Intel® Processor 3,0 GHz (i7-1185G7) | 16 GB | 512 GB SSD hard drive | 1000 Mbit network card, wireless LAN 802.11 ac | contact manufacturer | contact manufacturer | contact manufacturer |  | Yes | Yes, up to date graphic user interface, intuitive operation via touchscreen and an adaptable organ specific image processing. It also includes a diagnostic module standard and the multimedia radiographic positioning guide. | Yes; contact manufacturer for details | Yes COP (cognition optimised processing) including: automatic dead pixel correction automatic image area analysis automatic noise filter gridless imaging automatic grid line suppression automatic contrast optimisation intelligent brightness control | Yes | Yes; Qure.ai (optional) | Yes; contact manufacturer for details |  | contact manufacturer | Yes (optional) | Yes (optional) |  | FDA 510k / CE-IVDD / Aus TGA | Yes | Yes |  | Yes | Yes | Yes | Yes 2 years standard warranty | Yes | Yes |  | Yes, upon request Cost of 2-days of training, travel and accommodation to be added | contact manufacturer |  | contact manufacturer | No | contact manufacturer
OR TECHNOLOGY | Germany | https://www.oehm-rehbein.de/de | info@or-technology.com |  | P-100/20HB | Yes | 11.2 kg (X-ray generator only) | Alredy assembled |  | 1.6 kW | 40 - 110 kV (1 kV steps) | 0.4 - 50 mAs | 20 mA | contact manufacturer | Stationary, Tungsten Anode, 1.2 mm focal spot | 20 kHU | Yes; through Capacitor | Manual shutter with high brightness and integrated timer dual laser pointer min light field: ≤ 5 x 5 cm at 100cm max light field: ≤ 47 x 47 cm at 100cm auto timer: 30s | contact manufacturer | Yes; separate component. Various mounting solutions including wall mounts are available | At 65 cm SID, 35 cm x 35 cm X-ray field | Yes | 11.2 kg (including battery) | Battery, 130 - 260 VDC (automatic power adjustment) | 3 hours | 300 exposures | No | No | Yes | Wired, two-stage hand trigger switch and remote control hand switch | Temperature (without battery): -25 to 55°C Temperature (with battery): -20 to 33°C Humidity: 10 - 95% (non condensing) Atmospheric pressure: 500 - 1060 hPa | Temperature: 10 to 40 °C Humidity: 30 - 75 % (non condensing) Atmospheric pressure: 700 - 1060 hPa | Yes; Protection against electric shock class 1 type B under IEC 60601-1 standard |  | 35 x 43 cm | 3 sec | 100 µm | 4.3 lp/mm | 73.4% at 0 lp/mm | contact manufacturer | Battery (24 V) | AC (Battery charger including 2x batteries) | contact manufacturer for exposure numbers battery runtime of 8 hours | approx. 4 hours | No | Yes | wired or wireless | 3.0 kg | Yes dicomPACS DX-R | Temperature: -20 to 55 °C Humidity: 5 - 95 % (non-condensing) Atmospheric pressure: 600 - 1060 hPa Altitude: max. 3000 meters | Temperature: -20 to 55°C Humidity: 5 - 95 % (non-condensing) Atmospheric pressure: 600 - 1060 hPa Altitude: max. 3000 meters | contact manufacturer | Yes • Mobile special stand for DR detector / CR cassette and grid (35x43cm) • Available for different cassette formats, please specify detector type (manufacturer model) when ordering • Very flexible positioning of cassette and grid (due to vertical, horizontal and swivel joints as well as integrated spirit level) • Weight: 20 kg |  | mini workstation | 24" touch screen monitor, 1920 x 1080 Pixel | English | Intel® Processor 3,2 GHz (i7-8700) | 16 GB | 256 GB SSD hard drive 2 x 2 TB hard drive mirrored (RAID 1) | 2x 1000 Mbit network card | contact manufacturer | contact manufacturer | contact manufacturer |  | Yes | Yes, up to date graphic user interface, intuitive operation via touchscreen and an adaptable organ specific image processing. It also includes a diagnostic module standard and the multimedia radiographic positioning guide. | Yes; contact manufacturer for details | Yes COP (cognition optimised processing) including: automatic dead pixel correction automatic image area analysis automatic noise filter gridless imaging automatic grid line suppression automatic contrast optimisation intelligent brightness control | Yes | Yes; Qure.ai (optional) | Yes; with carrying case |  | contact manufacturer | Yes (optional) | Yes (optional) |  | CE | Yes | contact manufacturer |  | Yes | Yes | Yes | Yes 3 years manufacturer warranty (hand trigger switch excluded) 12 months warranty on X-ray tube and battery | Yes | Yes |  | Yes, upon request Cost of 2-days of training, travel and accommodation to be added | contact manufacturer |  | contact manufacturer | No | contact manufacturer
OR TECHNOLOGY | Germany | https://www.oehm-rehbein.de/de | info@or-technology.com |  | Amadeo P-100/35HB | Yes | 14.2 kg (X-ray generator only) | Already assembled |  | 2.4 kW | 40 - 100 kV (1 kV steps) | 0.4 - 100 mAs | 35 mA | contact manufacturer | Stationary, Tungsten Anode, 1.2 mm focal spot | 20 kHU | Yes; through Capacitor | Manual shutter with high brightness and integrated timer dual laser pointer min light field 5 cm x 5 cm at 100 cm max light field 35 cm x 35 cm <t 100 cm Lamp: LED Auto timer 30 sec | contact manufacturer | Yes; separate component. Various mounting solutions including wall mounts are available | At 100 cm SID, ≤47 cm x 47 cm X-ray field | Yes | 14.2 kg (including battery) | Battery (Lithium Polymer) AC 110 - 220 V (Free-Voltage) for charging | 3 hours | 400 exposures | No | No | Yes | Wired, two-stage hand trigger switch and remote control hand switch | Temperature (without battery): -25 to 60°C Temperature (with battery): -20 to 33°C Humidity: 10 - 95% (non condensing) Atmosphere pressure: 500 - 1060 hPa | Temperature: 10 to 40 °C Humidity: 30 - 75 % (non condensing) Atmospheric pressure: 700 - 1060 hPa | Yes; Protection against electric shock class 1 type B under IEC 60601-1 standard |  | 35 x 43 cm | 3 sec | 100 µm | 4.3 lp/mm | 73.4% at 0 lp/mm | contact manufacturer | Battery (24 V) | AC (Battery charger including 2x batteries) | contact manufacturer for exposure numbers battery runtime of 8 hours | approx. 4 hours | No | Yes | wired or wireless | 3.0 kg | Yes dicomPACS DX-R | Temperature: -20 to 55 °C Humidity: 5 - 95 % (non-condensing) Atmospheric pressure: 600 - 1060 hPa Altitude: max. 3000 meters | Temperature: -20 to 55 °C Humidity: 5 - 95 % (non-condensing) Atmospheric pressure: 600 - 1060 hPa Altitude: max. 3000 meters | contact manufacturer | Yes • Mobile special stand for DR detector / CR cassette and grid (35x43cm) • Available for different cassette formats, please specify detector type (manufacturer model) when ordering • Very flexible positioning of cassette and grid (due to vertical, horizontal and swivel joints as well as integrated spirit level) • Weight: 20 kg |  | mini worksataion | 24" touch screen monitor, 1920 x 1080 Pixel | English | Intel® Processor 3,2 GHz (i7-8700) | 16 GB | 256 GB SSD hard drive 2 x 2 TB hard drive mirrored (RAID 1) | 2x 1000 Mbit network card | contact manufacturer | contact manufacturer | contact manufacturer |  | Yes | Yes, up to date graphic user interface, intuitive operation via touchscreen and an adaptable organ specific image processing. It also includes a diagnostic module standard and the multimedia radiographic positioning guide. | Yes; contact manufacturer for details | Yes COP (cognition optimised processing) including: automatic dead pixel correction automatic image area analysis automatic noise filter gridless imaging automatic grid line suppression automatic contrast optimisation intelligent brightness control | Yes | Yes; Qure.ai (optional) | Yes; with carrying case |  | contact manufacturer | Yes (optional) | Yes (optional) |  | CE | Yes | contact manufacturer |  | Yes | Yes | Yes | Yes 3 years manufacturer warranty (hand trigger switch excluded) 12 months warranty on X-ray tube | Yes | Yes |  | Yes, upon request Cost of 2-days of training, travel and accommodation to be added | contact manufacturer |  | contact manufacturer | No | contact manufacturer
OR TECHNOLOGY | Germany | https://www.oehm-rehbein.de/de | info@or-technology.com |  | Amadeo P-110/100H | Yes | 20.5 kg (X-ray generator only) | Already assembled |  | 5 kW | 40 - 110 kV (1 kV steps) | 0.1 - 100 mAs | 50 mA | contact manufacturer | Stationary, Tungsten Anode, 1.5 mm focal spot | 31 kHU | Yes; through Capacitor | Manual shutter with high brightness and integrated timer dual laser pointer Min light field : 5 cm x 5 cm at 100 cm Max. light field : 47 cm x 47 cm at 100 cm Lamp : 3.3 V, LED Auto timer : 30 sec | contact manufacturer | Yes; separate component. Various mounting solutions including wall mounts are available | 100 cm | Yes | Approx. 20.5 kg | AC, 100 - 240 V(Free-Voltage) Capacitor backup at limited capacity | no battery available , capacitor backup for limited use, capacitor charging while in use | Unlimited exposures wit AC up to 20 exposure while used on fully charged capacitor | Not applicable No battery available, capacitor charging while in use | Not applicable | Yes, for the capacitor | Wired, two-stage hand trigger switch and remote control hand switch | Temperature: -10 to 55 °C Humidity: 10 - 90% Atmospheric pressure: 700 - 1060 hPa If the unit was transported or stored at temperatures under or above the recommended temperature range, a warm up or cool down period must be observed prior to operation of the unit | Temperature: 10 to 30 °C Humidity: 50 - 70 % (non condensing) Atmospheric pressure: 700 - 1060 hPa | Yes; Protection against electric shock class 1 type B under IEC 60601-1 standard |  | 35 x 43 cm | 3 sec | 100 µm | 4.3 lp/mm | 73.4% at 0 lp/mm | contact manufacturer | Battery (24 V) | AC (Battery charger including 2x batteries) | contact manufacturer for exposure numbers battery runtime of 8 hours | approx. 4 hours | No | Yes | wireless | 3.0 kg | Yes dicomPACS DX-R | Temperature: -20 to 55 °C Humidity: 5 - 95 % (non-condensing) Atmospheric pressure: 600 - 1060 hPa Altitude: max. 3000 meters | Temperature: -20 to 55 °C Humidity: 5 - 95 % (non-condensing) Atmospheric pressure: 600 - 1060 hPa Altitude: max. 3000 meters | contact manufacturer | Yes • Mobile special stand for DR detector / CR cassette and grid (35x43cm) • Available for different cassette formats, please specify detector type (manufacturer model) when ordering • Very flexible positioning of cassette and grid (due to vertical, horizontal and swivel joints as well as integrated spirit level) • Weight: 20 kg |  | mini worksataion | 24" touch screen monitor, 1920 x 1080 Pixel | English | Intel® Processor 3,2 GHz (i7-8700) | 16 GB | 256 GB SSD hard drive 2 x 2 TB hard drive mirrored (RAID 1) | 2x 1000 Mbit network card | contact manufacturer | contact manufacturer | contact manufacturer |  | Yes | Yes, up to date graphic user interface, intuitive operation via touchscreen and an adaptable organ specific image processing. It also includes a diagnostic module standard and the multimedia radiographic positioning guide. | Yes; contact manufacturer for details | Yes COP (cognition optimised processing) including: automatic dead pixel correction automatic image area analysis automatic noise filter gridless imaging automatic grid line suppression automatic contrast optimisation intelligent brightness control | Yes | Yes; Qure.ai (optional) | Yes; with carrying case |  | contact manufacturer | Yes (optional) | Yes (optional) |  | CE | Yes | contact manufacturer |  | Yes | Yes | Yes | Yes 3 years manufacturer warranty (hand trigger switch excluded) 12 months warranty on X-ray tube | Yes | Yes |  | Yes, upon request Cost of 2-days of training, travel and accommodation to be added | contact manufacturer |  | contact manufacturer | No | contact manufacturer
OR TECHNOLOGY | Germany | https://www.oehm-rehbein.de/de | info@or-technology.com |  | Amadeo P-AX SIUI SR-8230S | Yes | 13 kg (X-ray generator only) | Already assembled |  | 5 kW | 40 - 110 kV (1 kV steps) | 0.1 - 100 mAs | 100 mA | contact manufacturer | Stationary, 0.6 mm and 1.8 mm dual focal spot | contact manufacturer | contact manufacturer | contact manufacturer | 3 mm Al | Yes; separate component. Collapsible, light-weight tripod | 100 cm | Yes | Approx. 13 kg (Net) Approx. 20 kg (Gross) | Battery | 2.5 hours | 200 exposures up to 8 hours of battery life | Yes, also the option to recharge the battery off the device | No | Yes | Wired, two-stage hand trigger switch and remote control hand switch | contact manufacturer | Temperature : 10 to 40°C Humidity: 30 -75% Atmospheric pressure : 700 - 1060 hPa | contact manufacturer |  | 35 x 43 cm | 3 sec | 100 µm | 4.3 lp/mm | 73.4% at 0 lp/mm | contact manufacturer | Battery (24 V) | AC (Battery charger including 2x batteries) | contact manufacturer for exposure numbers battery runtime of 8 hours | approx. 4 hours | No | Yes | wireless | 3.0 kg | Yes dicomPACS DX-R | Temperature: -20 to 55 °C Humidity: 5 - 95 % (non-condensing) Atmospheric pressure: 600 - 1060 hPa Altitude: max. 3000 meters | Temperature: -20 to 55 °C Humidity: 5 - 95 % (non-condensing) Atmospheric pressure: 600 - 1060 hPa Altitude: max. 3000 meters | contact manufacturer | Yes • Mobile special stand for DR detector / CR cassette and grid (35x43cm) • Available for different cassette formats, please specify detector type (manufacturer model) when ordering • Very flexible positioning of cassette and grid (due to vertical, horizontal and swivel joints as well as integrated spirit level) • Weight: 20 kg |  | mini worksataion | 24" touch screen monitor, 1920 x 1080 Pixel | English | Intel® Processor 3,2 GHz (i7-8700) | 16 GB | 256 GB SSD hard drive 2 x 2 TB hard drive mirrored (RAID 1) | 2x 1000 Mbit network card | contact manufacturer | contact manufacturer | contact manufacturer |  | Yes | Yes, up to date graphic user interface, intuitive operation via touchscreen and an adaptable organ specific image processing. It also includes a diagnostic module standard and the multimedia radiographic positioning guide. | Yes; contact manufacturer for details | Yes COP (cognition optimised processing) including: automatic dead pixel correction automatic image area analysis automatic noise filter gridless imaging automatic grid line suppression automatic contrast optimisation intelligent brightness control | Yes | Yes; Qure.ai (optional) | Yes; with carrying case |  | contact manufacturer | Yes (optional) | Yes (optional) |  | contact manufaturer | Yes | Yes |  | Yes | Yes | Yes | Yes 1-year manufacturer's warranty including X-ray tube and battery | Yes | Yes |  | Yes, upon request Cost of 2-days of training, travel and accommodation to be added | contact manufacturer |  | contact manufacturer | No | contact manufacturer
OTOM | South Korea | https://nooka-otom.com/eng/page/?pid=medical_2&set=true | contact@theotom.com |  | MINE ALNU | Yes | 3 kg (X-ray generator only) | Already assembled |  | 0.24 kW | 40-80 kV | 0.3-3.9 mAs | 3 mA (Fixed) | 0.1-1.3 sec | Stationary, 0.4mm focal spot | contact manufacturer | Yes; auto-programmed waiting time (in seconds) after exposure ranges as follows: 0.1 to 0.2 seconds exposure, wait for 10 seconds; 0.2 to 0.5 seconds exposure, wait for 20 seconds; 0.6 to 1.3 seconds exposure, wait for 30 seconds. For safety reasons, if there is no activity for 120 seconds, the device will automatically enter power-saving mode. | Fixed, square collimator with LED light | ≥ 2.8 mm Al | Yes; separate component weight: 31 Kg Size: (L)870mm X (W)730mm X (H)2000mm Convenience of movement 5 fixed wheels | At 45 cm SID 21 x 21 cm X-ray field | contact manufacturer | 2.8kg | Battery | 2 hours | 100 exposures | No | No | Yes | Exposure switch on device and wireless remote control | Temperature: 15 to 40°C Humidity: 5 - 85% Atmospheric pressure: 760 - 1060 hPa | Temperature: - 40 to 70°C Humidity: 5 - 95% Atmospheric pressure: 760 - 1060 hPa Altitude: Less than 2000 m | contact manufacturer |  | 43 x 43 cm | Wired 4.5 sec Wireless 5 sec | 140 μm | 3.57 lp/mm | 84% at 0.5 lp/mm, 31% at 2 lp/mm 62% at 1 lp/mm, 15% at 3 lp/mm | contact manufacturer | Battery | AC Adapter Input: 100 - 240 VAC, Output: 1 9V/3.42 A | contact manufacturer | contact manufacturer | contact manufacturer | Yes; 1 extra battery | wireless | 1.3 kg | Yes O-Viewr AI | Temperature: -20 to 60 °C Humidity: Less than 90% (non- condensing) | Temperature: 10 to 35°C Humidity: Less than 80% (non-condensing) | IP68 for dust and water-resistance contact manufacturer for details | Yes Simple M Stand (Manual up/down) Dimension: 70(L) x 48(W) x 187.2(H) cm 19kg |  | Laptop | contact manufacturer | English (Windows dependent) | i5, i7 | 8 GB | 256 GB SSD (can be extended) | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer |  | Yes | Yes | Yes; contact manufacturer for details | Yes clip, zoom, magnifier, invert, rotate, flip, annotations, measurements, digital collimation, etc.), image view, detail enhancement and noise suppression, tissue equalization, alphanumeric annotation of images | Yes | Yes O-View AI OTOM self developed AI software | Yes; contact manufacturer for details |  | contact manufacturer | Yes | Yes |  | CE, FDA 510k | ISO13485 | IEC60601 |  | Yes | Yes | Yes | Yes 1 year warranty | No | Yes |  | Yes | Yes |  | Yes Standard Warranty Duration: 1 Year, batteries 6 months Warranty extension available: 1 - 3 years | Yes, in India | contact manufacturer
PROGNOSYS | India | https://prognosysmedical.com/mobile/ | marketing@prognosysmedical.com |  | PRORAD ATLAS | Yes | 20 kg | 5 to 10 minutes |  | 0.14 kW | 70 kV (Fixed) | 0.02 - 2.6 mAs | 2 mA (Fixed) | 0.01 - 1.3 sec | Stationary, 0.4 mm focal spot | 6.3 kHU | Yes; error message appears | Cone collimator with cross hair laser light | 1.5 mm Al | Yes; separate component. High quality Tripod stand included | 130 cm | No | 2.8 kg | Battery, 11.1 V (Lithium-ion Polymer) | 4 hours | 100 exposures | No | No | Yes | Wired (2 m hand switch) and remote control hand switch | Temperature: 10 to 30°C Humidity : 30 - 75% (non-condensing). Atmospheric pressure: 700 - 1060 hPa | Temperature: -10 to 55°C Humidity : 20 - 90% (non-condensing) Atmospheric pressure: 500 - 1060 hPa | Yes; physical shock and drop tested from 1.5m height |  | 35 x 43 cm | 3 sec | 140 µm | 3.5 lp/mm | 65% at 0 lp/mm | 16 bit | Battery | contact manufacturer | contact manufacturer for exposure numbers battery runtime of 8 hours | 3 hours | No | Yes | wireless | 2.5 kg | Yes console software | contact manufacturer | contact manufacturer | contact manufacturer | Yes, tripod stand with detector holder |  | contact manufacturer | contact manufacturer | English | i5 | 8 GB /16 GB | 1 TB | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer |  | Yes | Yes | contact manufacturer | Yes; contact manufacturer for details | Yes | Yes; contact manufacturer for details | Yes; contact manufacturer for details |  | contact manufacturer | contact manufacturer | contact manufacturer |  | contact manufacturer | contact manufacturer | contact manufacturer |  | Yes | contact manufacturer | contact manufacturer | Yes 1 year Standard warranty | contact manufacturer | contact manufacturer |  | contact manufacturer | contact manufacturer |  | contact manufacturer | contact manufacturer | contact manufacturer
RADISEN | South Korea | https://radisentech.com/ | Contact FIND for contact info |  | SmartPDX | Yes | 62 kg | less than 5 minutes |  | 5 kW | 50 - 110 kV | 0.4 - 100 mAs | 50 mA | 0.004 - 5 sec | Stationary, 1.8 mm focal spot | 42 kHU | Yes,the machie is automatically programmed | Manual, shutter type | 3.0 mm Al | Yes; folding stand with automatic height adjustment device | 100 cm and 180 cm (72 inches) | Yes; folding stand can adjust X-ray angle | 12 kg | Battery AC (110 - 120 V) for charging | < 2 hours | 999 exposures | Yes | Yes | Yes | Wired (2 m hand switch) | Temperature: -25 to 60°C Humidity: 10 - 95% | Temperature: 10 to 40°C Humidity : 30 - 75% | contact manufacturer |  | 35 x 43 cm | < 3 sec | 140 µm | 3.4 lp/mm | 57% at 1 lp/mm 40% at 2 lp/mm 10% at 3.5 lp/mm | contact manufacturer | Battery (12.6 V), DC (15 V) | contact manufacturer | contact manufacturer | 3 hours | Yes | Yes | wired or wireless | 3.2 kg | Yes | Temperature: -15 to 44°C Humidity: 10 - 95% Atmospheric pressure: 500 - 1060 hPa | Temperature: -15 to 44°C Humidity: 10 - 95% Atmospheric pressure: 500 - 1060 hPa | contact manufacturer | contact manufacturer |  | Laptop | 1920 x 1080 pixel | English | i7 | >8 GB | >500 GB | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer |  | Yes | Yes | Yes; Manual | Yes; contact manufacturer for details | Yes | Yes; contact manufacturer for details | Yes; contact manufacturer for details |  | contact manufacturer | contact manufacturer | contact manufacturer |  | Yes, contact manufacturer for further details | Yes | contact manufacturer |  | Yes | Yes | contact manufacturer | Yes 1 year standard warranty | contact manufacturer | contact manufacturer |  | Yes | Yes |  | Yes Standard warranty : X-ray generator 1 year, Detector 2 years, Battery 6 months (with conditions) Extension available | Yes, in Philippines | contact manufacturer
RadmediX | United States | https://radmedix.com/solutions/acuity-pdr/ | info@radmedix.com |  | Acuity PDR/PDR+ | Yes | 65 kg (X-ray generator only) | contact manufacturer |  | 2 kW / 5 kW options | 40 - 100 kV / 40 - 125 kV options | 50 mAs max / 200 mAs max options | 20 - 25 mA / 20 - 100 mA options | contact manufacturer | Stationary, 1.2 mm focal spot | contact manufacturer | contact manufacturer | Manual, rotating shutters LED/Laser light | contact manufacturer | Yes; built into the device (weight 11.7 kg) | At 65 cm SID ≤ 35cm x 35cm X-ray field | contact manufacturer | 11.8 kg | Battery, 29.6 VDC AC, 100-240 V, 50/60Hz for charging | contact manufacturer | 300 exposures | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer |  | 43 x 35 cm 43 x 43 cm | contact manufacturer | 140 μm | 3.4 lp/mm | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | 3.4 - 4.5 kg | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer |  | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer |  | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer |  | contact manufacturer | contact manufacturer | contact manufacturer |  | FDA 510k | contact manufacturer | contact manufacturer |  | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer |  | contact manufacturer | contact manufacturer |  | contact manufacturer | contact manufacturer | contact manufacturer
Rayence | South Korea | https://www.rayenceusa.com/mobile | information@rayenceusa.com |  | GODPX | Yes | contact manufacturer | contact manufacturer |  | 2.4 kW (at 80kV) | 40 -100 kV | 0.4 -100 mAs | 35 mA | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | AC, 110 - 220 V, single phase, 50 - 60 Hz | Not applicable | unlimited exposures on AC | Not applicable | Not applicable | Not applicable | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer |  | 43 x 35 cm | ≤ 5 sec | 140 μm | 5.5 lp/mm | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | wired | 3 kg | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer |  | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer |  | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer |  | contact manufacturer | contact manufacturer | contact manufacturer |  | contact manufacturer | contact manufacturer | contact manufacturer |  | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer |  | contact manufacturer | contact manufacturer |  | contact manufacturer | contact manufacturer | contact manufacturer
REMEDI | South Korea | https://remedihc.com/product-handheld-x-ray/ | sales@remedihc.com |  | REMEX-KA6 | Yes | 2.4 kg (X-ray generator only) | less than 5 minutes |  | 0.42 kW | 40 - 70 kV | 0.12 - 12 mAs | 6 mA | 0.06 - 2.0 sec | Stationary,Tungsten anode , 0.4 mm focal spot | contact manufacturer | Yes; when the device reaches 60°C, an integrated sensor detects the temperature and automatically shuts down the exposure system until the temperature falls below 60°C | contact manufacturer | ≥ 1.6 mm AI | Yes; separate component. 4 types 1) Manual height adjustable tripod stand 2) Manual height adjustable tripod stand with wheel 3) Automatic height adjustable stand 4) Carbon generator mount on manual height adjustable tripod stand | 100 cm | contact manufacturer | 2.4 kg | Battery, 22.2 VDC, 1800mAh AC (100-240 V, 50/60 Hz, 1.5A) for charging | 1.1 hours (approx) | 290 exposures (approx) | No | No | Yes | Wired (5 m hand switch) Tablet or notebook program available for remote exposure | Temperature: -40 to 70°C Humidity: 5 - 90% (non-condensing) Atmospheric pressure: 760 - 1060 hPa | Temperature: 10 to 40°C Humidity: 10 - 85 % (non-condensing) Atmospheric pressure: 800 - 1060 hPa Altitude: Less than 2000 m | contact manufacturer |  | 43 x 43 cm 35 x 43 cm | 3-5 sec (approx.) | 140 µm | 3.6 lp/mm | 62 % at 0.5 lp/mm 46 % at 1 lp/mm 40% at 2 lp/mm 27% at 3 lp/mm | 14 bit/16 bit | Battery | AC (100 - 240V) | contact manufacturer | approx. 4 hours | Yes | Yes (1 extra battery) | wired or wireless | 2.6 kg | Yes Zview | Temperature: -20 to 70°C Humidity: 10 - 95% (non-condensing) | Temperature: 10 to 35°C Humidity: 20 - 75 % (non-condensing) | contact manufacturer | Yes 3 types, size control detector mount with chin rest and 1) manual height adjustable tripod stand 2) Automatic height adjustable stand 3) Carbon detector mount on manual height adjustable tripod stand |  | Laptop, 1.75 kg (LG ultrabook, Windows 11) or Tablet, 0.6kg (Dicle Tab iWORK 20, Windows 11) | Laptop: 39.6cm (15.6") / 1920 x 1080 Tablet: 25.6cm (10.1") / 1920 x 1200 | English, Korean, French, Spanish, etc. | Core i5 Processor 1240P (1.7GHz) | 8 GB / DDR4 or 16 GB/DDR4 | 256 GB NVMe | Wired/Wireless data transfer | contact manufacturer | contact manufacturer | Contact manufacturer |  | Yes | Yes | contact manufacturer | Yes; auto processing | Yes | Yes; contact manufacturer for details | Yes; Backpack, Detector Case, Carrier Avaiable |  | 4 - 209 µSv/h Measuring conditions: at 70kV, 6mA, 1m from Focal Spot) | Yes | Yes |  | CE-IVDD / FDA 510k / Aus ARTG / KFDA and others, contact manufacturer for comprehensive details | Yes | Yes |  | Yes | Yes | Yes | Yes Various warranty programs for generatorand detector (18months, 36 months, 60 months) Laptop,batteries, AI software (up to 1 year) | No | Yes |  | No very easy to use, half day training will suffice Additional remote training, video guide can be provided upon request | Optional; depends on contract Adequate instruction are provided in the user manual Additional onsite training, video guides can be provided upon request |  | contact manufacturer | Yes, experience with TB programs in several countries, contact manufacturer for details | contact manufacturer
SEDECAL | Spain | https://www.sedecal.com/en/products/dragon-lw/ | info@sedecal.com |  | Dragon LW | Yes | 75.5 kg (X-ray generator only) | contact manufacturer |  | 4 kW | 40 - 125 kV ( 1 kV steps) | 0.1 - 250 mAs | 100 mA | 0.001 - 10 sec | Stationary, 0.5 mm and 1.8 mm dual focal spot | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | Yes; Generator Integrated with Stand, wheeled base | contact manufacturer | contact manufacturer | contact manufacturer | AC, 100 - 240 V, 50 - 60 Hz | Not applicable | unlimited exposures on AC | Not applicable | Not applicable | Not applicable | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer |  | contact manufacturer | contact manufacturer | 140 μm | 3.57 lp/mm | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer |  | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer |  | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer |  | contact manufacturer | contact manufacturer | contact manufacturer |  | FDA 510k | contact manufacturer | contact manufacturer |  | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer | contact manufacturer |  | contact manufacturer | contact manufacturer |  | contact manufacturer | contact manufacturer | contact manufacturer
VERTICE | South Africa | https://vertice.co.za/ | andries.vorster@vertice.co.za |  | LTE Nano TRB90 | Yes | 35 kg | 10 minutes |  | 1.35 kW | 40-90 kV (2kV steps) | 0.2 - 20 mAs | 15 mA | 0.01 - 1.0 sec | Stationary, Tungsten anode, 0.8 mm focal spot | 10 kHU | Yes; The unit is programmed with overload protection, an error message is generated during overload condition and the system will not expose | Manual multi leaf with LED light, dual laser pointer | > 2.5 mmAl | Yes; separate component | Adjustable SID, 120cm - 160cm | Yes; - LTE Portable CXR stand that offers electical movement and tracking of generator and DR panel, includes SID adjustment and fixed patient positioning. - Tri-pod stand setup will require use of internal laser pointer built into generator. | 7.3 kg | Battery, 57.6 VDC (Lithium ion) | Approx. 2 - 3 hours | > 200 exposures | No | Yes, upon request | Yes | Wired exposure switch | Temperature: -20°C to 50°C Humidity: max 85% Atmospheric pressure: 500 - 1060hPa | Temperature: 10°C to 40°C Humidity: 30-75% Atmospheric pressure: 700 - 1060hPa | contact manufacturer |  | 36 x 43 cm 43 x 43 cm | < 3 sec | 140 μm | 3.6 lp/mm |  | 16 bit | Battery, Li-ion (or via tether cable) | AC (battery charger) | ± 250 | 3 hours | Yes | Yes (1 extra battery) | wired or wireless | 2.9kg (35x43cm) / 3.4 kg (43x43cm) | Yes DRTech eConsole | Temperature : -15 to 55 °C Humidity : 10 - 90% Atmospheric pressure : 500 - 1060 hPa | Temperature : 10 to 35 °C Humidity : 20 - 85% Atmospheric pressure : 700 - 1060 hPa | IP54 | Yes, tri-pod detector stand as standard (other stand and holders available on request) |  | Laptop | 15.6' OLED 2.8k (2880 x 1620) integrated display | Multi language | Intel | 32 GB | 1 TB SSD | via DICOM or LAN or WAN | contact manufacturer | contact manufacturer | contact manufacturer |  | Yes | Yes | Yes | Yes; contact manufacturer for details | Yes | Yes; Qure qXR or other | Yes; Backpack and hard cases |  | contact manufacturer | Yes, 0.35mm Pb apron, and 0.35mm Pb thyroid shield | Yes, 0.35mm Pb half apron |  | CE certification | Yes | Yes |  | Yes | Yes | Yes | Yes 1 year warranty (2 or 3 year extension available) | Yes | Yes |  | Yes; training provided by qualified X-Ray technician and radiographer. Training included for Southern Africa, and available for African region and beyond. | Yes; case-by-case discussion based on customer needs |  | Yes extended warranty available with option of full SLA. Helpdesk support. | Yes, currently providing hardware for TB screening in over 22 African countries. Based out of South Africa with regional support partnners in Lesotho, Botswana, Zimbabwe, Mozambique, eSwatini, Zambia, Ivory Coast, DRC, Malawi, Tanzania. | contact manufacturer
VERTICE | South Africa | https://vertice.co.za/ | andries.vorster@vertice.co.za |  | LTE Nano PXM20 | Yes | 45 kg | 10 minutes |  | 1.6 kW | 40-100 kV (1kV steps) | 0.4 - 50 mAs | 20 mA | 0.01 - 1.0 sec | Stationary, Tungsten anode, 1.2 mm focal spot | 20 kHU | Yes; The unit is programmed with overload protection, an error message is generated during overload condition and the system will not expose | Manual multi leaf with LED light, dual laser pointer | > 2.5 mmAl | Yes; separate component | Adjustable SID, 120cm - 160cm | Yes; - LTE Portable CXR stand that offers electical movement and tracking of generator and DR panel, includes SID adjustment and fixed patient positioning. - Tri-pod stand setup will require use of internal laser pointer built into generator. | 11.5 kg | Battery, 14.8 VDC (Lithium Polymer) | 3.5 hours | Up to 200 exposures (at 90kV) | No | Yes, upon request | Yes | Wired exposure switch | Temperature: -20°C to 50°C Humidity: max 85% Atmospheric pressure: 500 - 1060hPa | Temperature: 10°C to 40°C Humidity: 30-75% Atmospheric pressure: 700 - 1060hPa | contact manufacturer |  | 36 x 43 cm 43 x 43 cm | < 3 sec | 140 μm | 3.6 lp/ mm |  | 16 bit | Battery, Li-ion (or via tether cable) | AC (battery charger) | ± 250 | 3 hours | Yes | Yes (1 extra battery) | wired or wireless | 2.9kg (35x43cm) / 3.4 kg (43x43cm) | Yes DRTech eConsole | Temperature : -15 to 55 °C Humidity : 10 - 90% Atmospheric pressure : 500 - 1060 hPa | Temperature : 10 to 35 °C Humidity : 20 - 85% Atmospheric pressure : 700 - 1060 hPa | IP54 | Yes, tri-pod detector stand as standard (other stand and holders available on request) |  | Laptop | 15.6' OLED 2.8k (2880 x 1620) integrated display | Multi language | Intel | 32 GB | 1 TB SSD | via DICOM or LAN or WAN | contact manufacturer | contact manufacturer | contact manufacturer |  | Yes | Yes | Yes | Yes; contact manufacturer for details | Yes | Yes; Qure qXR or other | Yes; Backpack and hard cases |  | contact manufacturer | Yes, 0.35mmPb apron, and 0.35mmPb thyroid shield | Yes, 0.35mm Pb half apron |  | CE certification | Yes | Yes |  | Yes | Yes | Yes | Yes 1 year warranty (2 or 3 year extension available) | Yes | Yes |  | Yes; training provided by qualified X-Ray technician and radiographer. Training included for Southern Africa, and available for African region and beyond. | Yes; case-by-case discussion based on customer needs |  | Yes extended warranty available with option of full SLA. Helpdesk support. | Yes, currently providing hardware for TB screening in over 22 African countries. Based out of South Africa with regional support partnners in Lesotho, Botswana, Zimbabwe, Mozambique, eSwatini, Zambia, Ivory Coast, DRC, Malawi, Tanzania. | contact manufacturer